## Cover Sheet
Analytical Cover Sheet
QAPP: | In Vitro TK Assessments of PFAS (D-CED-0031343)
Analysis Set ID: | 2020_PFAS_HepCl_AMIDES_ALK | *Hepatic clearance assay of PFAS acids
Analytes: | DTXSID2060965, DTXSID60400587, DTXSID70366226, DTXSID80310730 | PFAS
Dates Prepared: | 2020-12-02 | HepCl
Assay Preparer: | EKB, MJP
Dates Prepared: | 12/4/2020, 12/14/2020 | HepCl_CC
Dates of MS Run: | 2/10/2020, 3/30/2020
Assay Preparer: | ALK
Report Preparer: | ALK
Sample Type: | Standards, QCs, Unknown Analyte Samples | CC1-CC15
Internal Standard (conc) | 13C6-4-Nitrotoluene [lot SDFK-011] (25 pg/uL) | Amides; Ref
MDL Calculation |  |  |  | References:
|  |  |  | 40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
eLOQ Definition | Lowest curve concentration at which precision checks can be quantitated against curve ±30%
QC Samples and Acceptance Criteria:
QC Type | Description | Acceptance Criteria | Type
QC Sample Type | Frequency | Acceptance Criteria | QC Metric | Corrective Action
(Analytical, TK Assay, or Both) |  |  | Assessed | if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | r2≥0.98; | Linearity | Rerun samples and curve
7-pt curve min.
Curve Check (Analytical) | 2-3 samples per run | 75-125 % | Accuracy | Flag in comments
Instrument Blanks (Analytical) | 1 every 6-10 injections | <½ LOD of instrument method | Specificity | Flag in comments, assess and resolve
Technical Replicate assessment | Every assay sample | 75-125% | Precision | Flag in comments
Reference Chemical(s) (Assay) | 1 per assay batch | RSD = ±20% of historical or published values | Reproducibility | Flag in comments
Analyte-IS Matching
Analyte | Name | Sample ID | IS | Reference Compound | Avg. MW
DTXSID2060965 | Heptafluorobutyramide | 908 | 4NT-13C6 | * | 213.055
DTXSID60400587 | Nonafluoropentanamide | 909 |  |  | 263.063
DTXSID70366226 | Perfluoropentanamide | 916 |  |  | 245.072
DTXSID80310730 | Octafluoroadipamide | 923 |  |  | 288.097
*Preliminary method optimization work on these amides displayed similar clearance rates as seen here. Moreover, this hepatocyte lot has displayed high reproducibility across other assay runs
Sample ID Key
CC | Calibration Curve
QC | Quality Check
HT | Heat-treated Inactivated
Review of Data Timeline
Person | Action
EKB | Assay conducted 12/2/20, CC prepped 12/4/20
MJP | QC points prepped 12/14/20
ALK | Samples prepped & run 2/9/21, 2/22/21
MSC | Analytical data generated 2/10/20 & 3/30/20
ALK | QC Report completed 6/3/21
Notes
Concentrations listed are pre-crash concentrations in the assay
Subtraction of background, abiotic clearance was performed using the “No cell” negative controls. Heat inactivated hepatocytes were included in the assay but the resulting data indicates that, despite only 1% viability, metabolism was still occurring.
Hepatic Clearance |  |  |  |  | Bkgd-Sub Clearance (µL/min/106 hepatocytes)
Assay Date | MS Date | Species | Compound | Conc. (nM) |  | Flags
2020-12-02 | 2021-02-10 | Human | 908 | 1000 | 19.95 | accuracy in samples & inactivated & media controls
| 2021-02-10 |  | 909 | 1000 | 26.026 | unstable; T240 a & b <LOQ, so LOQ/sqrt(2) used
| 2021-03-30 |  | 923 |  | 2.954 | accuracy in samples & inactivated & media controls
| 2021-02-10 |  | 916 |  | 14.682 | accuracy in samples & inactivated & media controls
Stability in Williams' Media
Assay Date | MS Date | Compound | % of Time 0 at 60 min Avg | % of Time 0 at 60 min SD | Flags
2021-03-15 | 2021-03-15 | 923 | 94.35% | 37.18% | -
|  | 4NT | 47.29% | 19.05% | -
Quantitative Limits
DTXSID | Name | Sample ID | MDL (nM) | eLOQ (nM)
DTXSID2060965 | Heptafluorobutyramide | 908 | 1.81 | 11
DTXSID60400587 | Nonafluoropentanamide | 909 | 1.05 | 4.4
DTXSID70366226 | Perfluoropentanamide | 916 | 1.9 | 7.1
DTXSID80310730 | Octafluoroadipamide | 923 | 0.87 | 2.8

## Hep Clearance Calcs
Hepatocyte metabolic stabilty assay - Time course data and Clint calculations from Prism
Linear Regression Results
Analyte | Name | Sample ID | IS | Reference Compound | % Abiotic Loss | % Abiotic Loss StDev | Concentration (nM) | Best-Fit Values | [8 empty] | No Cells BkgdCl |  |  |  | HT BkgdCl |  |  |  | Bkgd-Sub Clearance (µL/min/106 hepatocytes)
[5 empty] | (Mean +/- SD) | (Mean +/- SD) |  | Slope | r2 | Sy.x | F | DFn | DFd | P value | Half Life | Clearance | T-test pvalue | Slope | Half Life | Clearance | T-test pvalue | Slope | Half Life | Clearance
DTXSID2060965 | Heptafluorobutyramide | 908 | 4NT-13C6 | 4NT | 28.76% | 47.77% | 1000 | -0.009975 | 0.9138 | 0.2676 | 169.6 | 1 | 16 | <0.0001 | 69.4737 | 19.95 | 0.4168 | NS | 0 | 0 | 0.3522 | NS | NS | 0 | 19.95
DTXSID60400587 | Nonafluoropentanamide** | 909 |  |  | 65.72% | 2.379% |  | -0.01747 | 0.8596 | 0.6167 | 97.97 | 1 | 16 | <0.0001 | 39.668 | 34.94 | 0.0006 | -0.004457 | 155.4858 | 8.914 | 0.0492 | -0.007311 | 94.7887 | 14.622 | 26.026
DTXSID80310730 | Octafluoroadipamide | 923 |  |  | 24.2% | 11.41% |  | -0.002646 | 0.3297 | 0.3296 | 7.87 | 1 | 16 | 0.0127 | 261.9048 | 5.292 | 0.0206 | -0.001169 | 592.8144 | 2.338 | 0.0005 | -0.001514 | 457.7279 | 3.028 | 2.954
DTXSID70366226 | Perfluoropentanamide | 916 |  |  | -52.8% | 113.5% |  | -0.007341 | 0.7174 | 0.4025 | 40.61 | 1 | 16 | <0.0001 | 94.4013 | 14.682 | 0.1809 | NS | 0 | 0 | 0.2122 | NS | NS | 0 | 14.682
*Preliminary method optimization work on these amides displayed similar clearance rates as seen here. Moreover, this hepatocyte lot has displayed high reproducibility across other assay runs
**unstable- 66% loss between T0 & T240 media only controls
Clint calcs:
k
0

## Hep Data for Prism
3/23/20 LC Runs of 3/4/20 & 3/6/20 Hepatocytes
Calcs
|  |  | 908 | [6 empty] | ln Transformed for Prism Entry- Clint Calcs on Hep Clear Data Tab
2020-12-02 | Human | 1uM | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 908 | T=0 act | 988.2968 | 123.9641 |  | 797.2442 | 23.2217 | 1 uM | Y1 | Y2 | Y3
|  |  | 908 | T=0 act | 784.7758 | 98.4361 |  |  |  | 0 | 6.896 | 6.6654 | 6.4276
indicates <LOQ, so LOD/sqrt(2) used |  |  | 908 | T=0 act | 618.6601 | 77.5998 |  |  |  | 15 | 6.311 | 6.8639 | 6.5722
|  |  | 908 | T=15 act | 550.6184 | 74.3187 |  | 740.8881 | 27.5988 | 30 | 6.2135 | 6.3299 | 6.3776
|  |  | 908 | T=15 act | 957.0957 | 129.1822 |  |  |  | 60 | 6.0332 | 5.7483 | 5.8549
|  |  | 908 | T=15 act | 714.95 | 96.4991 |  |  |  | 120 | 5.3074 | 5.5849 | 5.4239
|  |  | 908 | T=30 act | 499.4708 | 90.861 |  | 549.7087 | 8.2982 | 240 | 3.4785 | 4.5106 | 4.6317
|  |  | 908 | T=30 act | 561.1179 | 102.0755
|  |  | 908 | T=30 act | 588.5372 | 107.0635
|  |  | 908 | T=60 act | 417.0331 | 115.8824 |  | 359.8761 | 14.6009
|  |  | 908 | T=60 act | 313.6682 | 87.1601
|  |  | 908 | T=60 act | 348.927 | 96.9575
|  |  | 908 | T=120 act | 201.8328 | 87.1271 |  | 231.6532 | 14.0473
|  |  | 908 | T=120 act | 266.3618 | 114.983
|  |  | 908 | T=120 act | 226.7651 | 97.8899
|  |  | 908 | T=240 act | 32.412 | 43.0104 |  | 75.3584 | 49.9632
|  |  | 908 | T=240 act | 90.9721 | 120.7193
|  |  | 908 | T=240 act | 102.6911 | 136.2703
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD
|  |  | 908 | T=0 Inac | 235.4541 | 38.2703 |  | 615.2399 | 69.4934 | Heat Inac
|  |  | 908 | T=0 Inac | 531.9624 | 86.4642 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 908 | T=0 Inac | 1078.3033 | 175.2655 |  |  |  | 0 | 5.4615 | 6.2766 | 6.9831
|  |  | 908 | T=0 Media | 931.0556 | 147.488 |  | 631.2754 | 46.4629 | 240 | 5.2264 | 5.811 | 6.0887
|  |  | 908 | T=0 Media | 344.8974 | 54.635
|  |  | 908 | T=0 Media | 617.8731 | 97.877 |  |  |  | Media
|  |  | 908 | T=240 Inac | 186.1134 | 58.1046 |  | 320.3076 | 39.9356 | 1 uM | Y1 | Y2 | Y3
|  |  | 908 | T=240 Inac | 333.9572 | 104.2614 |  |  |  | 0 | 6.8363 | 5.8432 | 6.4263
|  |  | 908 | T=240 Inac | 440.8523 | 137.634 |  |  |  | 240 | 6.0143 | 6.1637 | 6.1414
|  |  | 908 | T=240 Media | 409.2301 | 90.9993 | 56.05% | 449.7069 | 7.8813
|  |  | 908 | T=240 Media | 475.1814 | 105.6647 | -37.77% |  |  | % Abiotic Loss | 28.76% | % Abiotic Loss StDev | 47.77%
|  |  | 908 | T=240 Media | 464.7091 | 103.336 | 24.79%
909
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 909 | T=0 act | 390.1618 | 75.9283 |  | 513.8556 | 25.5846 | 1 uM | Y1 | Y2 | Y3
|  |  | 909 | T=0 act | 651.9171 | 126.8678 |  |  |  | 0 | 5.9666 | 6.4799 | 6.2136
|  |  | 909 | T=0 act | 499.488 | 97.204 |  |  |  | 15 | 5.3701 | 4.84 | 5.2634
|  |  | 909 | T=15 act | 214.8822 | 120.6071 |  | 178.1671 | 25.8562 | 30 | 5.2587 | 5.0361 | 5.2591
|  |  | 909 | T=15 act | 126.475 | 70.9867 |  |  |  | 60 | 4.3134 | 4.3373 | 4.0226
|  |  | 909 | T=15 act | 193.144 | 108.4061 |  |  |  | 120 | 3.1222 | 3.065 | 2.5916
|  |  | 909 | T=30 act | 192.2322 | 107.1088 |  | 179.4737 | 12.3535 | 240 | 1.135 | 1.135 | 3.0732
|  |  | 909 | T=30 act | 153.8725 | 85.7354
|  |  | 909 | T=30 act | 192.3166 | 107.1558
|  |  | 909 | T=60 act | 74.6949 | 108.2314 |  | 69.0141 | 16.578
|  |  | 909 | T=60 act | 76.5031 | 110.8515
|  |  | 909 | T=60 act | 55.8443 | 80.9172
|  |  | 909 | T=120 act | 22.6965 | 118.4557 |  | 19.1603 | 26.4642
|  |  | 909 | T=120 act | 21.4336 | 111.8646
|  |  | 909 | T=120 act | 13.3509 | 69.6797
|  |  | 909 | T=240 act | 3.1113 | 33.5338 |  | 9.278 | 115.1229
|  |  | 909 | T=240 act | 3.1113 | 33.5338
|  |  | 909 | T=240 act | 21.6115 | 232.9324
|  |  | Compound | Type | Conc. (nM) | Precision | %Abiotic Loss | Avg Conc. (nM) | RSD
|  |  | 909 | T=0 Inac | 235.4541 | 38.2703 |  | 615.2399 | 69.4934 | Heat Inac
|  |  | 909 | T=0 Inac | 531.9624 | 86.4642 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 909 | T=0 Inac | 1078.3033 | 175.2655 |  |  |  | 0 | 5.4615 | 6.2766 | 6.9831
|  |  | 909 | T=0 Media | 364.9945 | 104.499 |  | 349.2805 | 13.4267 | 240 | 3.597 | 4.8183 | 5.0422
|  |  | 909 | T=0 Media | 386.3024 | 110.5994
|  |  | 909 | T=0 Media | 296.5448 | 84.9016 |  |  |  | Media
|  |  | 909 | T=240 Inac | 36.488 | 34.744 |  | 105.0197 | 58.4161 | 1 uM | Y1 | Y2 | Y3
|  |  | 909 | T=240 Inac | 123.7554 | 117.8402 |  |  |  | 0 | 5.8999 | 5.9566 | 5.6922
|  |  | 909 | T=240 Inac | 154.8157 | 147.4158 |  |  |  | 240 | 4.8885 | 4.809 | 4.6423
|  |  | 909 | T=240 Media | 132.756 | 110.892 | 63.63% | 119.7164 | 12.2823
|  |  | 909 | T=240 Media | 122.6139 | 102.4203 | 68.26% |  |  | % Abiotic Loss | 65.72% | % Abiotic Loss StDev | 2.379%
|  |  | 909 | T=240 Media | 103.7794 | 86.6877 | 65%
923
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 923 | T=0 act | 263.3388 | 114.3531 |  | 230.2856 | 12.5474 | 1 uM | Y1 | Y2 | Y3
|  |  | 923 | T=0 act | 217.7007 | 94.5351 |  |  |  | 0 | 5.5734 | 5.3831 | 5.3462
|  |  | 923 | T=0 act | 209.8174 | 91.1118 |  |  |  | 15 | 5.6765 | 5.4163 | 6.0269
|  |  | 923 | T=15 act | 291.9118 | 94.0242 |  | 310.4644 | 30.9401 | 30 | 5.9711 | 5.7497 | 5.8187
|  |  | 923 | T=15 act | 225.036 | 72.4837 |  |  |  | 60 | 6.0803 | 6.1268 | 5.7111
|  |  | 923 | T=15 act | 414.4454 | 133.4921 |  |  |  | 120 | 5.6815 | 5.8254 | 5.6278
|  |  | 923 | T=30 act | 391.931 | 112.7812 |  | 347.5144 | 11.5304 | 240 | 5.444 | 4.683 | 4.82
|  |  | 923 | T=30 act | 314.083 | 90.3799
|  |  | 923 | T=30 act | 336.5291 | 96.8389
|  |  | 923 | T=60 act | 437.1749 | 109.5382 |  | 399.1073 | 21.1887
|  |  | 923 | T=60 act | 457.9483 | 114.7432
|  |  | 923 | T=60 act | 302.1986 | 75.7186
|  |  | 923 | T=120 act | 293.3945 | 96.6982 |  | 303.4127 | 10.4122
|  |  | 923 | T=120 act | 338.7989 | 111.6627
|  |  | 923 | T=120 act | 278.0447 | 91.6391
|  |  | 923 | T=240 act | 231.3611 | 149.7739 |  | 154.4735 | 43.4103
|  |  | 923 | T=240 act | 108.0965 | 69.9773
|  |  | 923 | T=240 act | 123.9631 | 80.2487
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD
|  |  | 923 | T=0 Inac | 209.1715 | 100.67 |  | 207.7793 | 0.63118 | Heat Inac
|  |  | 923 | T=0 Inac | 207.5994 | 99.9134 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 923 | T=0 Inac | 206.5671 | 99.4166 |  |  |  | 0 | 5.3432 | 5.3356 | 5.3306
|  |  | 923 | T=0 Media | 214.7093 | 94.1687 |  | 228.0049 | 5.7015 | 240 | 5.0362 | 4.9188 | 4.9641
|  |  | 923 | T=0 Media | 240.6869 | 105.5622
|  |  | 923 | T=0 Media | 228.6185 | 100.2691 |  |  |  | Media
|  |  | 923 | T=240 Inac | 153.8854 | 106.3946 |  | 144.6365 | 5.9569 | 1 uM | Y1 | Y2 | Y3
|  |  | 923 | T=240 Inac | 136.838 | 94.6082 |  |  |  | 0 | 5.3693 | 5.4835 | 5.4321
|  |  | 923 | T=240 Inac | 143.1861 | 98.9972 |  |  |  | 240 | 5.1657 | 5.022 | 5.2554
|  |  | 923 | T=240 Media | 175.1641 | 111.7928 | 18.42% | 172.8256 | 11.6013
|  |  | 923 | T=240 Media | 151.7089 | 96.8233 | 36.97% |  |  | % Abiotic Loss | 24.2% | % Abiotic Loss StDev | 11.41%
|  |  | 923 | T=240 Media | 191.6037 | 122.2849 | 16.19%
916
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 916 | T=0 act | 266.4145 | 47.4477 |  | 561.4909 | 49.3501 | 1 uM | Y1 | Y2 | Y3
|  |  | 916 | T=0 act | 601.8854 | 107.1942 |  |  |  | 0 | 5.5851 | 6.4001 | 6.7046
|  |  | 916 | T=0 act | 816.1729 | 145.3582 |  |  |  | 15 | 6.2359 | 5.352 | 6.5976
|  |  | 916 | T=15 act | 510.7829 | 105.3049 |  | 485.0516 | 54.0368 | 30 | 6.5052 | 6.1058 | 6.1249
|  |  | 916 | T=15 act | 211.0287 | 43.5064 |  |  |  | 60 | 6.6295 | 5.9892 | 6.0804
|  |  | 916 | T=15 act | 733.3433 | 151.1887 |  |  |  | 120 | 5.4414 | 5.752 | 5.4039
|  |  | 916 | T=30 act | 668.5907 | 127.4182 |  | 524.7216 | 23.7591 | 240 | 4.8208 | 4.2328 | 4.4896
|  |  | 916 | T=30 act | 448.468 | 85.4678
|  |  | 916 | T=30 act | 457.1061 | 87.114
|  |  | 916 | T=60 act | 757.0915 | 142.5413 |  | 531.1383 | 37.0162
|  |  | 916 | T=60 act | 399.0999 | 75.1405
|  |  | 916 | T=60 act | 437.2235 | 82.3182
|  |  | 916 | T=120 act | 230.7708 | 118.8476 |  | 255.953 | 19.9837
|  |  | 916 | T=120 act | 314.8108 | 162.1285
|  |  | 916 | T=120 act | 222.2775 | 114.4736
|  |  | 916 | T=240 act | 124.0698 | 131.9618 |  | 94.0195 | 29.6865
|  |  | 916 | T=240 act | 68.9066 | 73.2897
|  |  | 916 | T=240 act | 89.0821 | 94.7486
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD
|  |  | 916 | T=0 Inac | 770.4552 | 112.9468 |  | 682.1401 | 18.9379 | Heat Inac
|  |  | 916 | T=0 Inac | 742.0912 | 108.7887 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 916 | T=0 Inac | 533.8737 | 78.2645 |  |  |  | 0 | 6.647 | 6.6095 | 6.2802
|  |  | 916 | T=0 Media | 370.3071 | 77.8272 |  | 423.8739 | 53.5143 | 240 | 5.2033 | 5.9381 | 6.567
|  |  | 916 | T=0 Media | 672.6962 | 141.3801
|  |  | 916 | T=0 Media | 228.6185 | 100.2691 |  |  |  | Media
|  |  | 916 | T=240 Inac | 181.8669 | 42.8825 |  | 424.1048 | 63.079 | 1 uM | Y1 | Y2 | Y3
|  |  | 916 | T=240 Inac | 379.2173 | 89.416 |  |  |  | 0 | 5.9143 | 6.5113 | 5.4321
|  |  | 916 | T=240 Inac | 711.23 | 167.7015 |  |  |  | 240 | 6.501 | 6.3412 | 6.565
|  |  | 916 | T=240 Media | 665.7928 | 102.7957 | -79.79% | 647.6858 | 11.2512
|  |  | 916 | T=240 Media | 567.4671 | 87.6146 | 15.64% |  |  | % Abiotic Loss | -52.8% | % Abiotic Loss StDev | 113.5%
|  |  | 916 | T=240 Media | 709.7973 | 109.5898 | -210.5%

## Assay Summary
12/2 Assay Layout | 12/7 CC Setup
*CC crashed 1:4, so dilutions & sample had to be adjusted to account for being 2x as dilute as samples

## MDL
908
Sample | [6 empty] | 908 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Final Conc.(nM) | Accuracy | Area
| HC_Amide CC 1 |  | 7010_1_051121010.D | Cal | 1 | 2021-05-11 16:53 | 9.2526 | 3.3369 | 119.1751 | 7668.0446
| HC_Amide CC 1 |  | 7010_1_051121011.D | Cal | 1 | 2021-05-11 17:26 | 9.2404 | 2.6535 | 94.7691 | 6134.4968
| HC_Amide CC 1 |  | 7010_1_051121012.D | Cal | 1 | 2021-05-11 17:59 | 9.2404 | 3.0603 | 109.2952 | 6930.6569
| HC_Amide CC 1 |  | 7010_1_051121013.D | Cal | 1 | 2021-05-11 18:32 | 9.2404 | 3.1177 | 111.3461 | 7223.6093
| HC_Amide CC 1 |  | 7010_1_051121014.D | Cal | 1 | 2021-05-11 19:06 | 9.2282 | 2.5627 | 91.5263 | 5577.6599
| HC_Amide CC 1 |  | 7010_1_051121015.D | Cal | 1 | 2021-05-11 19:39 | 9.2282 | 4.1389 | 147.8193 | 9034.0715
| HC_Amide CC 1 |  | 7010_1_051121016.D | Cal | 1 | 2021-05-11 20:12 | 9.216 | 3.8111 | 136.1093 | 8098.5725
[7 empty] | sd | 0.576
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  | MDL (nM) | 1.81
909
Sample | [6 empty] | 909 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Final Conc. | Accuracy | Area
| HC_Amide CC 1 |  | 7010_1_051121010.D | Cal | 1 | 2021-05-11 16:53 | 9.3372 | 2.8009 | 100.0317 | 1877.2423
| HC_Amide CC 1 |  | 7010_1_051121011.D | Cal | 1 | 2021-05-11 17:26 | 9.325 | 2.3055 | 82.3405 | 1573.4277
| HC_Amide CC 1 |  | 7010_1_051121012.D | Cal | 1 | 2021-05-11 17:59 | 9.3372 | 2.5292 | 90.3273 | 1681.7319
| HC_Amide CC 1 |  | 7010_1_051121013.D | Cal | 1 | 2021-05-11 18:32 | 9.3494 | 3.2724 | 116.8727 | 2185.8992
| HC_Amide CC 1 |  | 7010_1_051121014.D | Cal | 1 | 2021-05-11 19:06 | 9.2517 | 2.7106 | 96.8081 | 1724.4887
| HC_Amide CC 1 |  | 7010_1_051121015.D | Cal | 1 | 2021-05-11 19:39 | 9.2151 | 2.8597 | 102.1319 | 1817.9053
| HC_Amide CC 1 |  | 7010_1_051121016.D | Cal | 1 | 2021-05-11 20:12 | 9.2029 | 3.1392 | 112.1148 | 1929.536
[7 empty] | sd | 0.334
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  | MDL (nM) | 1.05
916
Sample | [6 empty] | 916 Method |  | 916 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| HC_Amide CC 1 |  | 7010_1_051121010.D | Cal | 1 | 2021-05-11 16:53 | y = 0.391972 * x | 2.2204e-16 | 11.9459 | 3.5489 | 126.7473 | 7341.637
| HC_Amide CC 1 |  | 7010_1_051121011.D | Cal | 1 | 2021-05-11 17:26 | y = 0.391972 * x | 2.2204e-16 | 11.9459 | 2.8651 | 102.3247 | 6019.4044
| HC_Amide CC 1 |  | 7010_1_051121012.D | Cal | 1 | 2021-05-11 17:59 | y = 0.391972 * x | 2.2204e-16 | 11.9459 | 3.7215 | 132.912 | 7569.197
| HC_Amide CC 1 |  | 7010_1_051121013.D | Cal | 1 | 2021-05-11 18:32 | y = 0.391972 * x | 2.2204e-16 | 11.9357 | 2.3988 | 85.6721 | 5083.4335
| HC_Amide CC 1 |  | 7010_1_051121014.D | Cal | 1 | 2021-05-11 19:06 | y = 0.391972 * x | 2.2204e-16 | 11.9459 | 2.3099 | 82.4972 | 4603.7481
| HC_Amide CC 1 |  | 7010_1_051121015.D | Cal | 1 | 2021-05-11 19:39 | y = 0.391972 * x | 2.2204e-16 | 11.9459 | 2.2604 | 80.7273 | 4520.981
| HC_Amide CC 1 |  | 7010_1_051121016.D | Cal | 1 | 2021-05-11 20:12 | y = 0.391972 * x | 2.2204e-16 | 11.9357 | 2.5265 | 90.2334 | 4901.9371
[9 empty] | sd | 0.603
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.9
923
Sample | [6 empty] | 923 Method |  | 923 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| HC_Amides CC 1 |  | 7010_1_052521014.D | Cal | 1 | 2021-05-25 23:59 |  |  | 23.968 | 2.6573 | 94.9036 | 3296.2925
| HC_Amides CC 1 |  | 7010_1_052521015.D | Cal | 1 | 2021-05-26 00:32 |  |  | 23.9765 | 2.9903 | 106.7978 | 2886.9724
| HC_Amides CC 1 |  | 7010_1_052521016.D | Cal | 1 | 2021-05-26 01:06 |  |  | 23.9596 | 2.8006 | 100.0223 | 3104.1271
| HC_Amides CC 1 |  | 7010_1_052521017.D | Cal | 1 | 2021-05-26 01:39 |  |  | 23.9511 | 2.4591 | 87.8261 | 2853.8755
| HC_Amides CC 1 |  | 7010_1_052521018.D | Cal | 1 | 2021-05-26 02:12 |  |  | 23.9426 | 2.1747 | 77.6671 | 2307.1665
| HC_Amides CC 1 |  | 7010_1_052521019.D | Cal | 1 | 2021-05-26 02:45 |  |  | 23.9426 | 2.7639 | 98.7099 | 3129.8339
| HC_Amides CC 1 |  | 7010_1_052521020.D | Cal | 1 | 2021-05-26 03:19 |  |  | 23.9596 | 2.4116 | 86.1284 | 2668.2398
[9 empty] | sd | 0.276
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.87

## 908
Sample | [8 empty] | 908 Method | 908 Results | [6 empty] | Qualifier (165.9 -> 119.0) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | RT | Resp. | Ratio | MI
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021001.D | MatrixBlank |  | 2021-02-10 14:46 |  | 9.2521 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0142 | 471.2302 | 18.2619 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021002.D | MatrixBlank |  | 2021-02-10 15:19 |  | 9.3518 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0057 | 1098520.6023 | 16.617 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021003.D | Cal | 15 | 2021-02-10 15:52 | 2000 | 9.4016 | 1480437.8402 | FALSE | 1903.1623 |  | 1903.1623 | 95.1581 | 27.0112 | FALSE | 11.0057 | 2021923.1562 | 17.4464 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021004.D | Cal | 14 | 2021-02-10 16:25 | 1250 | 9.3945 | 1026364.1059 | FALSE | 1477.16 |  | 1477.16 | 118.1728 | 27.181 | FALSE | 11.0057 | 1806607.3565 | 17.0202 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021005.D | Cal | 13 | 2021-02-10 16:59 | 781 | 9.3945 | 598428.4981 | FALSE | 705.3826 |  | 705.3826 | 90.3179 | 28.3851 | FALSE | 11.0057 | 2201383.4929 | 17.6488 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021006.D | Cal | 12 | 2021-02-10 17:32 | 488 | 9.3945 | 345281.8904 | FALSE | 435.9324 |  | 435.9324 | 89.3304 | 30.3856 | FALSE | 11.0057 | 2046716.1773 | 17.3102 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021007.D | Cal | 11 | 2021-02-10 18:05 | 305 | 9.3945 | 230336.2358 | FALSE | 279.7584 |  | 279.7584 | 91.7241 | 27.9596 | FALSE | 11.0057 | 2113378.3928 | 17.4201 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021008.D | Cal | 10 | 2021-02-10 18:38 | 191 | 9.3945 | 130247.3745 | FALSE | 204.9539 |  | 204.9539 | 107.3057 | 24.1585 | FALSE | 11.0057 | 1619841.6805 | 16.9256 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021009.D | Cal | 9 | 2021-02-10 19:12 | 119 | 9.4016 | 92373.8072 | FALSE | 111.2347 |  | 111.2347 | 93.4746 | 28.9722 | FALSE | 11.0057 | 2070468.6065 | 17.1909 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021010.D | Cal | 8 | 2021-02-10 19:45 | 75 | 9.3945 | 50561.2101 | FALSE | 68.5839 |  | 68.5839 | 91.4452 | 20.949 | FALSE | 11.0015 | 1784585.9395 | 16.7087 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021011.D | Cal | 7 | 2021-02-10 20:18 | 47 | 9.3945 | 31697.8999 | FALSE | 61.3126 |  | 61.3126 | 130.4523 | 20.5374 | FALSE | 11.0015 | 1240301.514 | 17.0934 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021012.D | Cal | 6 | 2021-02-10 20:52 | 29 | 9.4016 | 18306.197 | FALSE | 32.2334 |  | 32.2334 | 111.1497 | 17.4231 | FALSE | 11.0057 | 1266186.0052 | 15.9545 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021013.D | Cal | 5 | 2021-02-10 21:25 | 18 | 9.3945 | 13994.6736 | FALSE | 27.1133 |  | 27.1133 | 150.6297 | 27.2579 | FALSE | 11.0057 | 1119241.1397 | 16.5974 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021014.D | Cal | 4 | 2021-02-10 21:58 | 11 | 9.3945 | 6019.4978 | FALSE | 7.8703 |  | 7.8703 | 71.5486 | 18.0599 | FALSE | 11.0015 | 1166503.0569 | 16.437 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021015.D | Cal | 3 | 2021-02-10 22:32 | 7.1 | 9.3945 | 6789.6918 | FALSE | 8.9525 |  | 8.9525 | 126.0922 | 16.4667 | FALSE | 11.0015 | 1218263.1222 | 16.2663 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021016.D | Cal | 2 | 2021-02-10 23:05 | 4.4 | 9.4016 | 4338.6181 | FALSE | 1.9887 |  | 1.9887 | 45.1979 | 14.8615 | FALSE | 11.0015 | 1487904.2374 | 16.5453 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021017.D | Cal | 1 | 2021-02-10 23:38 | 2.8 | 9.3945 | 3964.2475 | FALSE | 2.4634 |  | 2.4634 | 87.9801 | 13.4053 | FALSE | 11.0015 | 1279997.1288 | 16.3886 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021018.D | QC | 3 | 2021-02-11 00:11 | 7.1 | 9.3874 | 4281.8683 | FALSE | 2.2289 |  | 2.2289 | 31.3935 | 21.9335 | FALSE | 11.0015 | 1423685.6191 | 16.5971 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021019.D | MatrixBlank |  | 2021-02-11 00:45 |  | 9.3732 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0057 | 1295784.3796 | 16.5578 | FALSE
| Amide Hep120220 T0A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 21 | 7010_1_021021020.D | Sample |  | 2021-02-11 01:18 |  | 9.3945 | 371510.6748 | FALSE | 988.2968 |  | 988.2968 | 123.9641 | 20.18 | FALSE | 11.0057 | 976881.6283 | 16.4514 | FALSE
| Amide Hep120220 T0A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 22 | 7010_1_021021021.D | Sample |  | 2021-02-11 01:51 |  | 9.4016 | 360487.9993 | FALSE | 784.7758 |  | 784.7758 | 98.4361 | 20.3404 | FALSE | 11.0057 | 1192634.9348 | 16.585 | FALSE
| Amide Hep120220 T0A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 23 | 7010_1_021021022.D | Sample |  | 2021-02-11 02:24 |  | 9.3945 | 271918.5833 | FALSE | 618.6601 |  | 618.6601 | 77.5998 | 20.7464 | FALSE | 11.0015 | 1139502.7817 | 16.4602 | FALSE
| Amide Hep120220 T0A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 24 | 7010_1_021021023.D | Sample |  | 2021-02-11 02:58 |  | 9.3945 | 135074.1147 | FALSE | 235.4541 |  | 235.4541 | 38.2703 | 22.7359 | FALSE | 11.0015 | 1467241.8316 | 16.4948 | FALSE
| Amide Hep120220 T0A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 25 | 7010_1_021021024.D | Sample |  | 2021-02-11 03:31 |  | 9.3945 | 195097.8443 | FALSE | 531.9624 |  | 531.9624 | 62.7975 | 17.9392 | FALSE | 11.0015 | 949650.797 | 15.7524 | FALSE
| Amide Hep120220 T0A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 26 | 7010_1_021021025.D | Sample |  | 2021-02-11 04:04 |  | 9.3945 | 408463.6766 | FALSE | 1078.3033 |  | 1078.3033 | 137.4069 | 21.517 | FALSE | 11.0057 | 984620.2053 | 16.5316 | FALSE
| Amide Hep120220 T0A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 27 | 7010_1_021021026.D | Sample |  | 2021-02-11 04:37 |  | 9.3945 | 360506.3896 | FALSE | 931.0556 |  | 931.0556 | 147.488 | 19.2536 | FALSE | 11.0015 | 1006033.0484 | 16.7534 | FALSE
| Amide Hep120220 T0A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 28 | 7010_1_021021027.D | Sample |  | 2021-02-11 05:11 |  | 9.3945 | 202453.14 | FALSE | 344.8974 |  | 344.8974 | 54.635 | 21.2751 | FALSE | 11.0015 | 1512105.363 | 16.9221 | FALSE
| Amide Hep120220 T0A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 29 | 7010_1_021021028.D | Sample |  | 2021-02-11 05:44 |  | 9.3945 | 280817.7938 | FALSE | 617.8731 |  | 617.8731 | 97.877 | 24.58 | FALSE | 11.0015 | 1178283.8521 | 16.2921 | FALSE
| Amide Hep120220 TMidA1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 30 | 7010_1_021021029.D | Sample |  | 2021-02-11 06:17 |  | 9.3945 | 258245.4016 | FALSE | 550.6184 |  | 550.6184 | 74.3187 | 21.3395 | FALSE | 11.0015 | 1214804.9822 | 16.4161 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021030.D | QC | 2 | 2021-02-11 06:50 | 4.4 | 9.3945 | 5840.0053 | FALSE | 3.7294 |  | 3.7294 | 84.7584 | 12.3992 | FALSE | 11.0015 | 1631225.3389 | 16.7198 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021031.D | QC | 11 | 2021-02-11 07:24 | 305 | 9.3945 | 93105.2432 | FALSE | 198.1494 |  | 198.1494 | 64.967 | 21.5406 | FALSE | 11.0015 | 1196600.3583 | 16.9051 | FALSE
| Amide Hep120220 TMidA2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 31 | 7010_1_021021032.D | Sample |  | 2021-02-11 07:57 |  | 9.3945 | 303392.851 | FALSE | 957.0957 |  | 957.0957 | 129.1822 | 18.6961 | FALSE | 11.0015 | 823692.1843 | 16.6719 | FALSE
| Amide Hep120220 TMidA3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 32 | 7010_1_021021033.D | Sample |  | 2021-02-11 08:30 |  | 9.3945 | 264895.6553 | FALSE | 714.95 |  | 714.95 | 96.4991 | 19.268 | FALSE | 11.0015 | 961483.9696 | 16.0318 | FALSE
| Amide Hep120220 TMidA4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 33 | 7010_1_021021034.D | Sample |  | 2021-02-11 09:03 |  | 9.3945 | 251242.731 | FALSE | 499.4708 |  | 499.4708 | 90.861 | 19.1202 | FALSE | 10.9972 | 1301720.7842 | 16.4131 | FALSE
| Amide Hep120220 TMidA5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 34 | 7010_1_021021035.D | Sample |  | 2021-02-11 09:37 |  | 9.3945 | 295910.0135 | FALSE | 561.1179 |  | 561.1179 | 102.0755 | 23.5337 | FALSE | 11.0015 | 1366155.6963 | 16.6035 | FALSE
| Amide Hep120220 TMidA6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 35 | 7010_1_021021036.D | Sample |  | 2021-02-11 10:10 |  | 9.3945 | 301520.7687 | FALSE | 588.5372 |  | 588.5372 | 107.0635 | 25.6215 | FALSE | 11.0015 | 1327721.2641 | 16.5042 | FALSE
| Amide Hep120220 TMidA7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 36 | 7010_1_021021037.D | Sample |  | 2021-02-11 10:43 |  | 9.3945 | 269020.3991 | FALSE | 417.0331 |  | 417.0331 | 115.8824 | 24.7857 | FALSE | 10.9972 | 1666051.7762 | 16.5629 | FALSE
| Amide Hep120220 TMidA8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 37 | 7010_1_021021038.D | Sample |  | 2021-02-11 11:16 |  | 9.3945 | 198920.7637 | FALSE | 313.6682 |  | 313.6682 | 87.1601 | 24.4705 | FALSE | 10.9972 | 1631164.8763 | 16.5653 | FALSE
| Amide Hep120220 TMidA9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 38 | 7010_1_021021039.D | Sample |  | 2021-02-11 11:50 |  | 9.3945 | 175336.9088 | FALSE | 348.927 |  | 348.927 | 96.9575 | 24.852 | FALSE | 11.0015 | 1294679.5364 | 16.1028 | FALSE
| Amide Hep120220 TMidA10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 39 | 7010_1_021021040.D | Sample |  | 2021-02-11 12:23 |  | 9.3945 | 120908.6153 | FALSE | 201.8328 |  | 201.8328 | 87.1271 | 21.3804 | FALSE | 10.9972 | 1526331.4613 | 16.5777 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021021042.D | QC | 14 | 2021-02-11 13:29 | 1250 | 9.3874 | 371730.8143 | FALSE | 742.7127 |  | 742.7127 | 59.417 | 21.9899 | FALSE | 10.9972 | 1299106.2208 | 15.9401 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021043.D | QC | 15 | 2021-02-11 14:03 | 2000 | 9.3945 | 1245963.8947 | FALSE | 1958.2869 |  | 1958.2869 | 97.9143 | 19.1498 | FALSE | 10.9972 | 1653672.9933 | 16.2989 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021041.D | QC | 7 | 2021-02-11 12:56 | 47 | 9.3874 | 36134.8527 | FALSE | 53.6376 |  | 53.6376 | 114.1225 | 25.5678 | FALSE | 11.0015 | 1596975.2223 | 16.2858 | FALSE
| Amide Hep120220 TMidA11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 40 | 7010_1_021021044.D | Sample |  | 2021-02-11 14:36 |  | 9.3874 | 145330.0754 | FALSE | 266.3618 |  | 266.3618 | 114.983 | 24.6202 | FALSE | 10.9972 | 1399153.6007 | 16.5182 | FALSE
| Amide Hep120220 TMidA12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 41 | 7010_1_021021045.D | Sample |  | 2021-02-11 15:09 |  | 9.3945 | 123687.3901 | FALSE | 226.7651 |  | 226.7651 | 97.8899 | 21.6956 | FALSE | 10.9972 | 1393815.7663 | 16.3068 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021021046.D | Sample |  | 2021-02-11 15:42 |  | 9.3874 | 20351.9582 | FALSE | 38.9755 |  | 38.9755 | 110.1251 | 22.6425 | FALSE | 10.9972 | 1195003.7494 | 16.3011 | FALSE
| Amide Hep120220 T0B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 51 | 7010_1_021021057.D | Sample |  | 2021-02-11 16:20 |  | 9.3945 | 60052.904 | FALSE | 141.2884 |  | 141.2884 |  | 19.4194 | FALSE | 10.9972 | 1070637.3302 | 16.1526 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021058.D | Sample |  | 2021-02-11 16:53 |  | 9.423 | 49603.6084 | TRUE | 104.1636 |  | 104.1636 |  | 11.2747 | TRUE | 10.9972 | 1183429.5555 | 16.487 | FALSE
| Amide Hep120220 T0B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 53 | 7010_1_021021059.D | Sample |  | 2021-02-11 17:26 |  | 9.4159 | 29925.3853 | TRUE | 63.7994 |  | 63.7994 |  | 13.284 | TRUE | 10.9972 | 1129010.3124 | 16.5765 | FALSE
| Amide Hep120220 T0B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 54 | 7010_1_021021060.D | Sample |  | 2021-02-11 17:59 |  | 9.3945 | 75890.6942 | FALSE | 156.8259 |  | 156.8259 |  | 18.6866 | FALSE | 10.9972 | 1223561.0281 | 16.7504 | FALSE
| Amide Hep120220 T0B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 55 | 7010_1_021021061.D | Sample |  | 2021-02-11 18:32 |  | 9.3945 | 108004.7885 | FALSE | 247.1537 |  | 247.1537 |  | 20.9601 | FALSE | 10.9972 | 1118870.834 | 16.1968 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021062.D | Sample |  | 2021-02-11 19:06 |  | 9.3945 | 122753.2837 | FALSE | 217.2532 |  | 217.2532 |  | 28.5116 | FALSE | 10.9972 | 1442355.8709 | 16.7947 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021063.D | Sample |  | 2021-02-11 19:39 |  | 9.3945 | 98115.1235 | FALSE | 232.5501 |  | 232.5501 |  | 21.7815 | FALSE | 10.9972 | 1078776.6744 | 16.1556 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021064.D | QC | 3 | 2021-02-11 20:12 | 7.1 | 9.3874 | 3902.2743 | TRUE | 1.1402 |  | 1.1402 | 16.059 | 28.5135 | FALSE | 10.9972 | 1505421.7831 | 16.0836 | FALSE
| Amide Hep120220 T240B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 72 | 7010_1_021021065.D | Sample |  | 2021-02-11 20:49 |  | 9.3874 | 15836.091 | FALSE | 24.6619 |  | 24.6619 |  | 30.2517 | FALSE | 10.9972 | 1368939.113 | 16.2463 | FALSE
| Amide Hep120220 T240B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 73 | 7010_1_021021066.D | Sample |  | 2021-02-11 21:22 |  | 9.3874 | 12897.381 | TRUE | 17.5605 |  | 17.5605 |  | 19.5979 | FALSE | 10.9972 | 1455996.9248 | 16.771 | FALSE
| Amide Hep120220 T240B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 74 | 7010_1_021021067.D | Sample |  | 2021-02-11 21:55 |  | 9.3874 | 35365.9827 | FALSE | 63.8012 |  | 63.8012 |  | 25.6618 | FALSE | 10.993 | 1334235.7534 | 16.2788 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021068.D | Sample |  | 2021-02-11 22:28 |  | 9.3874 | 65446.8965 | FALSE | 132.521 |  | 132.521 |  | 21.6422 | FALSE | 10.9972 | 1240868.7448 | 15.9939 | FALSE
| Amide Hep120220 T240B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 76 | 7010_1_021021069.D | Sample |  | 2021-02-11 23:01 |  | 9.3874 | 52646.1765 | FALSE | 109.2231 |  | 109.2231 |  | 24.5032 | FALSE | 10.993 | 1200683.2968 | 15.7968 | FALSE
| Amide Hep120220 T240B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 77 | 7010_1_021021070.D | Sample |  | 2021-02-11 23:35 |  | 9.3874 | 52354.9119 | FALSE | 124.0503 |  | 124.0503 |  | 21.2386 | FALSE | 10.993 | 1057498.623 | 16.4488 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021071.D | Sample |  | 2021-02-12 00:08 |  | 9.423 | 43959.6501 | TRUE | 90.8791 |  | 90.8791 |  | 55.7944 | FALSE | 10.9972 | 1193151.4602 | 16.1777 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021072.D | QC | 6 | 2021-02-12 00:41 | 29 | 9.3874 | 22798.9903 | FALSE | 36.6851 |  | 36.6851 | 126.5002 | 22.4852 | FALSE | 10.993 | 1411114.9478 | 16.5335 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021073.D | QC | 12 | 2021-02-12 01:14 | 488 | 9.3874 | 293512.5228 | FALSE | 523.5866 |  | 523.5866 | 107.2923 | 24.1107 | FALSE | 10.993 | 1451334.565 | 16.7815 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021021074.D | Sample |  | 2021-02-12 01:48 |  | 9.3945 | 89111.8897 | FALSE | 143.8121 |  | 143.8121 |  | 23.5202 | FALSE | 10.9972 | 1561871.923 | 16.8299 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021021075.D | Sample |  | 2021-02-12 02:21 |  | 9.3874 | 122042.9657 | FALSE | 252.9752 |  | 252.9752 |  | 21.3637 | FALSE | 10.993 | 1235820.4763 | 17.052 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021021076.D | Sample |  | 2021-02-12 02:54 |  | 9.3945 | 66434.8611 | FALSE | 134.0515 |  | 134.0515 |  | 21.046 | FALSE | 10.9972 | 1245796.3823 | 17.1506 | FALSE
| Amide Hep120220 T240B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 78 | 7010_1_021021077.D | Sample |  | 2021-02-12 03:27 |  | 9.3874 | 45880.5344 | FALSE | 103.5926 |  | 103.5926 |  | 20.4285 | FALSE | 10.993 | 1100328.3884 | 16.9299 | FALSE
| Amide Hep120220 T240B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 79 | 7010_1_021021078.D | Sample |  | 2021-02-12 04:01 |  | 9.3874 | 42986.6923 | FALSE | 88.1892 |  | 88.1892 |  | 21.5993 | FALSE | 10.993 | 1200195.6816 | 17.4207 | FALSE
| Amide Hep120220 T240B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 80 | 7010_1_021021079.D | Sample |  | 2021-02-12 04:34 |  | 9.3874 | 36213.6683 | FALSE | 76.5913 |  | 76.5913 |  | 22.5223 | FALSE | 10.993 | 1153710.9073 | 17.2878 | FALSE
| Amide Hep120220 T0C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 81 | 7010_1_021021080.D | Sample |  | 2021-02-12 05:07 |  | 9.3945 | 12584.9066 | FALSE | 27.2633 |  | 27.2633 |  | 21.243 | FALSE | 10.993 | 1001908.3613 | 16.7672 | FALSE
| Amide Hep120220 T0C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 82 | 7010_1_021021081.D | Sample |  | 2021-02-12 05:40 |  | 9.3945 | 8230.6213 | FALSE | 16.5271 |  | 16.5271 |  | 19.7746 | FALSE | 10.993 | 972455.3572 | 16.7866 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021082.D | QC | 7 | 2021-02-12 06:14 | 47 | 9.3874 | 26310.723 | FALSE | 58.0591 |  | 58.0591 | 123.53 | 17.2699 | FALSE | 10.993 | 1082089.2888 | 16.6215 | FALSE
| Amide Hep120220 T0C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 83 | 7010_1_021021083.D | Sample |  | 2021-02-12 06:47 |  | 9.4016 | 5568.3968 | FALSE | 11.9067 |  | 11.9067 |  | 19.8463 | TRUE | 10.993 | 831032.5682 | 16.8694 | FALSE
| Amide Hep120220 T0C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 84 | 7010_1_021021084.D | Sample |  | 2021-02-12 07:20 |  | 9.3945 | 15435.7744 | FALSE | 33.5025 |  | 33.5025 |  | 14.6468 | FALSE | 10.993 | 1033039.7387 | 16.4451 | FALSE
| Amide Hep120220 T0C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 85 | 7010_1_021021085.D | Sample |  | 2021-02-12 07:53 |  | 9.3874 | 19018.633 | FALSE | 49.4683 |  | 49.4683 |  | 17.7271 | FALSE | 10.993 | 904113.176 | 16.4554 | FALSE
| Amide Hep120220 T0C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 86 | 7010_1_021021086.D | Sample |  | 2021-02-12 08:27 |  | 9.3874 | 7283.3588 | FALSE | 14.8313 |  | 14.8313 |  | 24.0991 | FALSE | 10.993 | 931779.9049 | 16.8305 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021087.D | QC | 10 | 2021-02-12 09:00 | 191 | 9.3874 | 94471.3205 | FALSE | 189.2047 |  | 189.2047 | 99.0601 | 18.6203 | FALSE | 10.993 | 1269923.3239 | 16.2772 | FALSE
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021088.D | MatrixBlank |  | 2021-02-12 09:33 |  | 9.4372 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0015 | 0 |  | FALSE
| Amide Hep120220 T0C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 87 | 7010_1_021021089.D | Sample |  | 2021-02-12 10:06 |  | 9.3874 | 25959.6403 | FALSE | 69.0339 |  | 69.0339 |  | 17.2087 | FALSE | 10.993 | 910738.7204 | 16.7235 | FALSE
| Amide Hep120220 T0C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 88 | 7010_1_021021090.D | Sample |  | 2021-02-12 10:39 |  | 9.3874 | 7233.8231 | FALSE | 6.8423 |  | 6.8423 |  | 16.9836 | FALSE | 10.993 | 1517162.4518 | 16.8494 | FALSE
| Amide Hep120220 T0C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 89 | 7010_1_021021091.D | Sample |  | 2021-02-12 11:21 |  | 9.3874 | 12296.4022 | FALSE | 18.207 |  | 18.207 |  | 19.3806 | FALSE | 10.993 | 1350538.834 | 16.6547 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021092.D | MatrixBlank |  | 2021-02-12 11:54 |  | 9.3803 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.993 | 1537915.2315 | 16.4698 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021093.D | QC | 15 | 2021-02-12 12:28 | 2000 | 9.3874 | 1266593.3187 | FALSE | 2266.0499 |  | 2266.0499 | 113.3025 | 22.4235 | FALSE | 10.993 | 1452077.3799 | 16.2193 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021094.D | QC | 14 | 2021-02-12 13:01 | 1250 | 9.3874 | 766022.887 | FALSE | 1332.9249 |  | 1332.9249 | 106.634 | 19.2745 | FALSE | 10.993 | 1494237.8182 | 16.518 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021095.D | QC | 13 | 2021-02-12 13:34 | 781 | 9.3874 | 488837.6221 | FALSE | 1017.6762 |  | 1017.6762 | 130.3043 | 21.5645 | FALSE | 10.9887 | 1248386.5293 | 16.8735 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021096.D | QC | 12 | 2021-02-12 14:07 | 488 | 9.3874 | 314110.0526 | FALSE | 611.776 |  | 611.776 | 125.3639 | 24.2438 | FALSE | 10.993 | 1331012.2455 | 16.383 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021097.D | QC | 11 | 2021-02-12 14:40 | 305 | 9.3874 | 191616.6337 | FALSE | 314.0282 |  | 314.0282 | 102.9601 | 21.4914 | FALSE | 10.993 | 1569499.7597 | 16.8251 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021098.D | QC | 10 | 2021-02-12 15:14 | 191 | 9.3874 | 110054.4004 | FALSE | 172.2375 |  | 172.2375 | 90.1767 | 21.4432 | FALSE | 10.993 | 1620578.5035 | 16.6787 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021099.D | QC | 9 | 2021-02-12 15:47 | 119 | 9.3874 | 72134.4263 | FALSE | 129.7943 |  | 129.7943 | 109.0708 | 24.0158 | FALSE | 10.993 | 1395207.6913 | 16.5331 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021100.D | QC | 8 | 2021-02-12 16:20 | 75 | 9.3874 | 44425.5608 | FALSE | 75.6367 |  | 75.6367 | 100.8489 | 21.0036 | FALSE | 10.993 | 1431953.6932 | 16.6109 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021101.D | QC | 7 | 2021-02-12 16:53 | 47 | 9.3874 | 32744.1773 | FALSE | 44.0042 |  | 44.0042 | 93.626 | 24.0096 | FALSE | 10.993 | 1727908.7341 | 16.5044 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021102.D | QC | 6 | 2021-02-12 17:27 | 29 | 9.3874 | 22005.2817 | FALSE | 27.7213 |  | 27.7213 | 95.5908 | 31.0497 | FALSE | 10.993 | 1727836.1962 | 16.412 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021103.D | QC | 5 | 2021-02-12 18:00 | 18 | 9.3874 | 14600.6854 | FALSE | 18.189 |  | 18.189 | 101.0499 | 16.2712 | FALSE | 10.993 | 1604833.3904 | 16.5686 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021104.D | QC | 4 | 2021-02-12 18:33 | 11 | 9.3874 | 8845.7488 | FALSE | 9.6949 |  | 9.6949 | 88.1352 | 30.425 | FALSE | 10.993 | 1510408.5477 | 16.4124 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021105.D | QC | 3 | 2021-02-12 19:07 | 7.1 | 9.3874 | 3358.8944 | FALSE | 0.51744 |  | 0.51744 | 7.2879 | 73.1449 | FALSE | 10.9887 | 1426574.3777 | 16.5389 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021106.D | QC | 2 | 2021-02-12 19:40 | 4.4 | 9.3874 | 3194.2785 | FALSE | 0 |  | 0 | 0 | 33.3597 | FALSE | 10.993 | 1660880.0813 | 16.5037 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021107.D | QC | 1 | 2021-02-12 20:13 | 2.8 | 9.3803 | 2065.2941 | FALSE | 0 |  | 0 | 0 | 38.1155 | FALSE | 10.9887 | 1514529.3891 | 16.5007 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021108.D | QC | 3 | 2021-02-12 20:46 | 7.1 | 9.3803 | 4291.1119 | FALSE | 0.55223 |  | 0.55223 | 7.7779 | 25.5966 | FALSE | 10.9887 | 1812284.3058 | 16.6834 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021109.D | MatrixBlank |  | 2021-02-12 21:20 |  | 9.366 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9887 | 1876327.1648 | 16.3154 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021110.D | Sample |  | 2021-02-12 21:53 |  | 9.3874 | 71455.0897 | FALSE | 138.7274 |  | 138.7274 |  | 22.5157 | FALSE | 10.9887 | 1296524.3149 | 16.4413 | FALSE
| Amide Hep120220 T240C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 102 | 7010_1_021021111.D | Sample |  | 2021-02-12 22:26 |  | 9.3803 | 5026.5732 | TRUE | 6.7578 |  | 6.7578 |  | 32.4441 | FALSE | 10.993 | 1061413.605 | 16.4233 | FALSE
| Amide Hep120220 T240C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 103 | 7010_1_021021112.D | Sample |  | 2021-02-12 22:59 |  | 9.3803 | 6942.5295 | FALSE | 0.6219 |  | 0.6219 |  | 53.9839 | FALSE | 10.993 | 2899512.1558 | 17.9619 | FALSE
| Amide Hep120220 T240C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 104 | 7010_1_021021113.D | Sample |  | 2021-02-12 23:33 |  | 9.3874 | 30210.7611 | FALSE | 20.7525 |  | 20.7525 |  | 46.8971 | FALSE | 10.993 | 2998218.9334 | 17.8928 | FALSE
| Amide Hep120220 T240C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 105 | 7010_1_021021114.D | Sample |  | 2021-02-13 00:06 |  | 9.3874 | 28059.9137 | FALSE | 25.2261 |  | 25.2261 |  | 49.7266 | FALSE | 10.9887 | 2381288.7606 | 17.9741 | FALSE
| Amide Hep120220 T240C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 106 | 7010_1_021021115.D | Sample |  | 2021-02-13 00:39 |  | 9.3803 | 20145.9166 | FALSE | 33.7791 |  | 33.7791 |  | 35.0229 | FALSE | 10.9887 | 1338808.5188 | 16.8042 | FALSE
| Amide Hep120220 T240C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 107 | 7010_1_021021116.D | Sample |  | 2021-02-13 01:12 |  | 9.3874 | 18818.0253 | FALSE | 28.6701 |  | 28.6701 |  | 26.0776 | FALSE | 10.9887 | 1436728.2402 | 16.4115 | FALSE
| Amide Hep120220 T240C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 108 | 7010_1_021021117.D | Sample |  | 2021-02-13 01:46 |  | 9.3874 | 15071.6416 | FALSE | 21.5996 |  | 21.5996 |  | 25.2876 | FALSE | 10.9887 | 1449263.6916 | 16.4455 | FALSE
| Amide Hep120220 T240C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 109 | 7010_1_021021118.D | Sample |  | 2021-02-13 02:19 |  | 9.3874 | 15943.4493 | FALSE | 28.8373 |  | 28.8373 |  | 27.0119 | FALSE | 10.9887 | 1211357.7857 | 16.428 | FALSE
| Amide Hep120220 T240C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 110 | 7010_1_021021119.D | Sample |  | 2021-02-13 02:52 |  | 9.3874 | 18121.392 | FALSE | 43.049 |  | 43.049 |  | 19.7864 | FALSE | 10.993 | 975022.2483 | 15.7866 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021120.D | QC | 11 | 2021-02-13 03:25 | 305 | 9.3874 | 98535.1472 | FALSE | 125.5805 |  | 125.5805 | 41.1739 | 21.7954 | FALSE | 10.9887 | 1967061.2052 | 16.5843 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021521001.D | QC | 11 | 2021-02-15 09:38 | 305 | 9.3875 | 183630.174 | FALSE | 435.4423 |  | 435.4423 | 142.768 | 23.3112 | FALSE | 10.9888 | 1089709.628 | 15.8899 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521002.D | QC | 15 | 2021-02-15 10:11 | 2000 | 9.3875 | 1502353.5677 | FALSE | 2234.948 |  | 2234.948 | 111.7474 | 25.769 | FALSE | 10.9888 | 1746420.9306 | 16.7178 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521003.D | QC | 15 | 2021-02-15 10:44 | 2000 | 9.3804 | 1284364.4559 | FALSE | 1770.4293 |  | 1770.4293 | 88.5215 | 18.0097 | FALSE | 10.9888 | 1885916.7109 | 16.4804 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021521004.D | Sample |  | 2021-02-15 11:17 |  | 9.3875 | 105886.5686 | FALSE | 188.5636 |  | 188.5636 |  | 24.5563 | FALSE | 10.9888 | 1428073.1821 | 16.0304 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021521005.D | Sample |  | 2021-02-15 11:50 |  | 9.3804 | 136037.7383 | FALSE | 270.406 |  | 270.406 |  | 23.9032 | FALSE | 10.9888 | 1290491.2495 | 15.8517 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021521006.D | Sample |  | 2021-02-15 12:24 |  | 9.3875 | 83891.476 | FALSE | 177.9461 |  | 177.9461 |  | 23.2839 | FALSE | 10.9888 | 1196894.7839 | 16.3618 | FALSE
| Amide Hep120220 TMidB1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 60 | 7010_1_021521007.D | Sample |  | 2021-02-15 12:57 |  | 9.3875 | 33663.3864 | FALSE | 59.603 |  | 59.603 |  | 24.5149 | FALSE | 10.9888 | 1351722.6553 | 16.3029 | FALSE
| Amide Hep120220 TMidB2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 61 | 7010_1_021521008.D | Sample |  | 2021-02-15 13:30 |  | 9.3875 | 96965.0553 | FALSE | 165.681 |  | 165.681 |  | 30.2432 | FALSE | 10.9888 | 1482501.358 | 15.7608 | FALSE
| Amide Hep120220 TMidB3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 62 | 7010_1_021521009.D | Sample |  | 2021-02-15 14:03 |  | 9.3875 | 30275.1112 | FALSE | 44.4441 |  | 44.4441 |  | 22.896 | FALSE | 10.9889 | 1583586.7699 | 16.5115 | FALSE
| Amide Hep120220 TMidB4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 63 | 7010_1_021521010.D | Sample |  | 2021-02-15 14:37 |  | 9.3875 | 17649.2687 | TRUE | 27.9629 |  | 27.9629 |  | 24.4244 | FALSE | 10.9889 | 1375845.0535 | 15.8883 | FALSE
| Amide Hep120220 TMidB5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 64 | 7010_1_021521011.D | Sample |  | 2021-02-15 15:10 |  | 9.3875 | 25347.3027 | FALSE | 38.868 |  | 38.868 |  | 19.8687 | FALSE | 10.9888 | 1491909.1765 | 16.5452 | FALSE
| Amide Hep120220 TMidB6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 65 | 7010_1_021521012.D | Sample |  | 2021-02-15 15:43 |  | 9.3875 | 15045.1802 | FALSE | 22.8067 |  | 22.8067 |  |  | FALSE | 10.9888 | 1385355.4716 | 16.0732 | FALSE
| Amide Hep120220 TMidB7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 66 | 7010_1_021521013.D | Sample |  | 2021-02-15 16:17 |  | 9.3804 | 11837.035 | FALSE | 15.001 |  | 15.001 |  | 25.0281 | FALSE | 10.9888 | 1501903.8251 | 15.9211 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021521014.D | QC | 4 | 2021-02-15 16:50 | 11 | 9.3804 | 8615.3858 | FALSE | 8.4771 |  | 8.4771 | 77.0643 | 18.7567 | FALSE | 10.9846 | 1597862.8245 | 16.0039 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521015.D | MatrixBlank |  | 2021-02-15 17:23 |  | 9.3733 | 3700.757 | TRUE | 0 |  | 0 |  | 16.7528 | TRUE | 10.9846 | 2065444.3364 | 16.3871 | FALSE
| Amide Hep120220 TMidB8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 67 | 7010_1_021521016.D | Sample |  | 2021-02-15 17:56 |  | 9.3875 | 25065.3309 | FALSE | 37.2746 |  | 37.2746 |  | 20.7486 | FALSE | 10.9846 | 1530077.5364 | 16.7428 | FALSE
| Amide Hep120220 TMidB9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 68 | 7010_1_021521017.D | Sample |  | 2021-02-15 18:30 |  | 9.3875 | 16229.9241 | FALSE | 21.5323 |  | 21.5323 |  | 22.7746 | FALSE | 10.9846 | 1564508.1695 | 16.4263 | FALSE
| Amide Hep120220 TMidB10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 69 | 7010_1_021521018.D | Sample |  | 2021-02-15 19:03 |  | 9.3875 | 11258.0104 | FALSE | 12.8858 |  | 12.8858 |  | 22.7512 | FALSE | 10.9889 | 1591418.6379 | 16.5783 | FALSE
| Amide Hep120220 TMidB11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 70 | 7010_1_021521019.D | Sample |  | 2021-02-15 19:36 |  | 9.3804 | 19216.5775 | FALSE | 26.7088 |  | 26.7088 |  | 31.1422 | FALSE | 10.9888 | 1556081.2329 | 16.1663 | FALSE
| Amide Hep120220 TMidB12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 71 | 7010_1_021521020.D | Sample |  | 2021-02-15 20:09 |  | 9.3804 | 42000.5926 | FALSE | 95.3852 |  | 95.3852 |  | 26.0626 | FALSE | 10.9846 | 1089125.1205 | 15.782 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021521021.D | QC | 6 | 2021-02-15 20:43 | 29 | 9.3804 | 22645.0116 | FALSE | 48.0345 |  | 48.0345 | 165.6363 | 27.2341 | FALSE | 10.9846 | 1105257.3151 | 15.5945 | FALSE
| Amide Hep120220 TMidC1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 90 | 7010_1_021521022.D | Sample |  | 2021-02-15 21:16 |  | 9.3804 | 9137.1303 | FALSE | 12.2495 |  | 12.2495 |  | 15.8572 | FALSE | 10.9846 | 1337523.5952 | 15.7728 | FALSE
| Amide Hep120220 TMidC2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 91 | 7010_1_021521023.D | Sample |  | 2021-02-15 21:49 |  | 9.3804 | 27608.7969 | FALSE | 49.9388 |  | 49.9388 |  | 21.2288 | FALSE | 10.9846 | 1301363.9414 | 15.9555 | FALSE
| Amide Hep120220 TMidC3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 92 | 7010_1_021521024.D | Sample |  | 2021-02-15 22:22 |  | 9.3875 | 4522.9615 | FALSE | 3.8204 |  | 3.8204 |  | 18.2171 | FALSE | 10.9846 | 1251213.9398 | 15.6121 | FALSE
| Amide Hep120220 TMidC4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 93 | 7010_1_021521025.D | Sample |  | 2021-02-15 22:56 |  | 9.3804 | 9127.3149 | FALSE | 18.0254 |  | 18.0254 |  | 12.5212 | FALSE | 10.9846 | 1010158.4132 | 15.6893 | FALSE
| Amide Hep120220 TMidC5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 94 | 7010_1_021521026.D | Sample |  | 2021-02-15 23:29 |  | 9.3875 | 3425.6535 | FALSE | 3.5887 |  | 3.5887 |  | 32.2466 | FALSE | 10.9846 | 971417.7653 | 15.9506 | FALSE
| Amide Hep120220 TMidC6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 95 | 7010_1_021521027.D | Sample |  | 2021-02-16 00:02 |  | 9.3733 | 1389.5847 | FALSE | 0 |  | 0 |  | 19.5471 | FALSE | 10.9846 | 1358371.3736 | 16.5259 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021521028.D | QC | 14 | 2021-02-16 00:35 | 1250 | 9.3804 | 318572.3174 | FALSE | 693.8945 |  | 693.8945 | 55.5116 | 19.287 | FALSE | 10.9846 | 1191185.6119 | 16.177 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521029.D | QC | 15 | 2021-02-16 01:09 | 2000 | 9.3804 | 1403576.5736 | FALSE | 2798.2494 |  | 2798.2494 | 139.9125 | 39.6875 | FALSE | 10.9846 | 1301790.9075 | 16.6447 | FALSE
| Amide Hep120220 TMidC7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 96 | 7010_1_021521030.D | Sample |  | 2021-02-16 01:42 |  | 9.3733 | 5001.5933 | TRUE | 3.8733 |  | 3.8733 |  | 33.2399 | FALSE | 10.9337 | 1375933.0656 | 16.6485 | TRUE
| Amide Hep120220 TMidC8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 97 | 7010_1_021521031.D | Sample |  | 2021-02-16 02:15 |  | 9.3804 | 5606.0302 | FALSE | 4.7246 |  | 4.7246 |  | 14.7237 | FALSE | 10.9888 | 1415696.2804 | 16.9998 | FALSE
| Amide Hep120220 TMidC9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 98 | 7010_1_021521032.D | Sample |  | 2021-02-16 02:48 |  | 9.3733 | 2827.6948 | FALSE | 0.64304 |  | 0.64304 |  | 36.7752 | FALSE | 10.9846 | 1177007.9129 | 16.2408 | FALSE
| Amide Hep120220 TMidC10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 99 | 7010_1_021521033.D | Sample |  | 2021-02-16 03:22 |  | 9.3733 | 5370.0966 | FALSE | 3.2935 |  | 3.2935 |  | 15.5543 | FALSE | 10.9846 | 1573056.5806 | 16.6627 | FALSE
| Amide Hep120220 TMidC11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 100 | 7010_1_021521034.D | Sample |  | 2021-02-16 03:55 |  | 9.3804 | 6716.8836 | FALSE | 6.1834 |  | 6.1834 |  | 16.602 | FALSE | 10.9846 | 1487167.1693 | 16.3383 | FALSE
| Amide Hep120220 TMidC12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 101 | 7010_1_021521035.D | Sample |  | 2021-02-16 04:28 |  | 9.3804 | 8921.3491 | FALSE | 13.4945 |  | 13.4945 |  | 11.764 | FALSE | 10.9846 | 1221022.527 | 16.1578 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021521036.D | Sample |  | 2021-02-16 05:01 |  | 9.3804 | 21149.3049 | FALSE | 32.412 |  | 32.412 | 43.0104 | 23.0804 | FALSE | 10.9846 | 1455972.8968 | 16.8817 | FALSE
| Amide Hep120220 T240A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 43 | 7010_1_021521037.D | Sample |  | 2021-02-16 05:35 |  | 9.3804 | 42059.3899 | FALSE | 90.9721 |  | 90.9721 | 120.7193 | 19.8434 | FALSE | 10.9846 | 1140475.4058 | 16.5723 | FALSE
| Amide Hep120220 T240A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 44 | 7010_1_021521038.D | Sample |  | 2021-02-16 06:08 |  | 9.3804 | 45659.6529 | FALSE | 102.6911 |  | 102.6911 | 136.2703 | 21.338 | FALSE | 10.9846 | 1104145.2921 | 16.3082 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021521039.D | QC | 9 | 2021-02-16 06:41 | 119 | 9.3804 | 64739.8993 | FALSE | 182.664 |  | 182.664 | 138.8444 | 22.706 | FALSE | 10.9846 | 900503.6595 | 16.8534 | FALSE
| Amide Hep120220 T240A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 45 | 7010_1_021521040.D | Sample |  | 2021-02-16 07:15 |  | 9.3804 | 78031.6649 | FALSE | 186.1134 |  | 186.1134 | 58.1046 | 18.7266 | FALSE | 10.9846 | 1065852.6822 | 16.3473 | FALSE
| Amide Hep120220 T240A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 46 | 7010_1_021521041.D | Sample |  | 2021-02-16 07:48 |  | 9.3804 | 123958.6112 | FALSE | 333.9572 |  | 333.9572 | 100.6651 | 18.7338 | FALSE | 10.9846 | 955690.2177 | 16.0574 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521044.D | MatrixBlank |  | 2021-02-16 09:27 |  | 9.4444 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9803 | 1813919.0893 | 16.2148 | FALSE
| Amide Hep120220 T240A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 49 | 7010_1_021521045.D | Sample |  | 2021-02-16 10:01 |  | 9.3804 | 192102.3511 | FALSE | 475.1814 |  | 475.1814 | 105.6647 | 19.9473 | FALSE | 10.9804 | 1045655.2677 | 16.0693 | FALSE
| Amide Hep120220 T240A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 50 | 7010_1_021521046.D | Sample |  | 2021-02-16 10:34 |  | 9.3804 | 195940.5313 | FALSE | 464.7091 |  | 464.7091 | 103.336 | 21.0681 | FALSE | 10.9804 | 1090324.8465 | 16.5326 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021521047.D | QC | 7 | 2021-02-16 11:07 | 47 | 9.3733 | 26391.318 | FALSE | 59.0681 |  | 59.0681 | 125.6767 | 19.785 | FALSE | 10.9804 | 1068480.751 | 16.333 | FALSE
| Amide Hep120220 T240A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 47 | 7010_1_021521042.D | Sample |  | 2021-02-16 08:21 |  | 9.3804 | 158908.54 | FALSE | 440.8523 |  | 440.8523 | 137.634 | 18.9999 | FALSE | 10.9846 | 931565.1466 | 16.0569 | FALSE
| Amide Hep120220 T240A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 48 | 7010_1_021521043.D | Sample |  | 2021-02-16 08:54 |  | 9.3804 | 241993.8198 | FALSE | 409.2301 |  | 409.2301 | 90.9993 | 17.9771 | FALSE | 10.9804 | 1526895.0041 | 15.9114 | FALSE

## 909
Sample | [8 empty] | 909 Method | 909 Results | [6 empty] | Qualifier (166.0 -> 119.0) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | RT | Resp. | Ratio | MI
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021001.D | MatrixBlank |  | 2021-02-10 14:46 |  | 9.4292 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0142 | 471.2302 | 18.2619 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021002.D | MatrixBlank |  | 2021-02-10 15:19 |  | 9.3509 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0057 | 1098520.6023 | 16.617 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021003.D | Cal | 15 | 2021-02-10 15:52 | 2000 | 9.4649 | 1200838.9293 | TRUE | 1919.2396 |  | 1919.2396 | 95.962 | 5.0338 | FALSE | 11.0057 | 2021923.1562 | 17.4464 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021004.D | Cal | 14 | 2021-02-10 16:25 | 1250 | 9.4649 | 874138.697 | TRUE | 1501.9957 |  | 1501.9957 | 120.1597 | 4.6321 | FALSE | 11.0057 | 1806607.3565 | 17.0202 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021005.D | Cal | 13 | 2021-02-10 16:59 | 781 | 9.4649 | 505815.6366 | TRUE | 660.7735 |  | 660.7735 | 84.6061 | 4.5574 | FALSE | 11.0057 | 2201383.4929 | 17.6488 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021006.D | Cal | 12 | 2021-02-10 17:32 | 488 | 9.4649 | 298835.2574 | TRUE | 410.9033 |  | 410.9033 | 84.2015 | 5.5071 | FALSE | 11.0057 | 2046716.1773 | 17.3102 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021007.D | Cal | 11 | 2021-02-10 18:05 | 305 | 9.4649 | 205610.7602 | TRUE | 270.5496 |  | 270.5496 | 88.7048 | 5.0727 | FALSE | 11.0057 | 2113378.3928 | 17.4201 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021008.D | Cal | 10 | 2021-02-10 18:38 | 191 | 9.4649 | 96688.5205 | TRUE | 164.5139 |  | 164.5139 | 86.1329 | 4.9678 | FALSE | 11.0057 | 1619841.6805 | 16.9256 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021009.D | Cal | 9 | 2021-02-10 19:12 | 119 | 9.4649 | 81684.8884 | TRUE | 108.2253 |  | 108.2253 | 90.9456 | 5.3099 | FALSE | 11.0057 | 2070468.6065 | 17.1909 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021010.D | Cal | 8 | 2021-02-10 19:45 | 75 | 9.4649 | 36178.6109 | TRUE | 55.368 |  | 55.368 | 73.824 | 4.1422 | FALSE | 11.0015 | 1784585.9395 | 16.7087 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021011.D | Cal | 7 | 2021-02-10 20:18 | 47 | 9.4649 | 28245.0189 | TRUE | 62.231 |  | 62.231 | 132.4063 | 4.9696 | FALSE | 11.0015 | 1240301.514 | 17.0934 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021012.D | Cal | 6 | 2021-02-10 20:52 | 29 | 9.4649 | 15506.9227 | FALSE | 33.3872 |  | 33.3872 | 115.1283 | 5.4724 | FALSE | 11.0057 | 1266186.0052 | 15.9545 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021013.D | Cal | 5 | 2021-02-10 21:25 | 18 | 9.4649 | 11193.8791 | TRUE | 27.2513 |  | 27.2513 | 151.3964 | 2.9959 | TRUE | 11.0057 | 1119241.1397 | 16.5974 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021014.D | Cal | 4 | 2021-02-10 21:58 | 11 | 9.4649 | 5035.7288 | FALSE | 11.7476 |  | 11.7476 | 106.7965 |  | FALSE | 11.0015 | 1166503.0569 | 16.437 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021015.D | Cal | 3 | 2021-02-10 22:32 | 7.1 | 9.4649 | 2985.8902 | TRUE | 6.6669 |  | 6.6669 | 93.8999 |  | FALSE | 11.0015 | 1218263.1222 | 16.2663 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021016.D | Cal | 2 | 2021-02-10 23:05 | 4.4 | 9.4578 | 1987.6446 | TRUE | 3.6328 |  | 3.6328 | 82.5644 |  | FALSE | 11.0015 | 1487904.2374 | 16.5453 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021017.D | Cal | 1 | 2021-02-10 23:38 | 2.8 | 9.4577 | 4529.6304 | FALSE | 9.6283 |  | 9.6283 | 343.8689 |  | FALSE | 11.0015 | 1279997.1288 | 16.3886 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021018.D | QC | 3 | 2021-02-11 00:11 | 7.1 | 9.4577 | 3091.3032 | FALSE | 5.906 |  | 5.906 | 83.1826 | 22.4166 | FALSE | 11.0015 | 1423685.6191 | 16.5971 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021019.D | MatrixBlank |  | 2021-02-11 00:45 |  | 9.3367 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 11.0057 | 1295784.3796 | 16.5578 | FALSE
| Amide Hep120220 T0A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 21 | 7010_1_021021020.D | Sample |  | 2021-02-11 01:18 |  | 9.4649 | 22891.9669 | FALSE | 64.047 |  | 64.047 |  |  | FALSE | 11.0057 | 976881.6283 | 16.4514 | FALSE
| Amide Hep120220 T0A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 22 | 7010_1_021021021.D | Sample |  | 2021-02-11 01:51 |  | 9.4649 | 13656.4203 | FALSE | 31.2107 |  | 31.2107 |  |  | FALSE | 11.0057 | 1192634.9348 | 16.585 | FALSE
| Amide Hep120220 T0A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 23 | 7010_1_021021022.D | Sample |  | 2021-02-11 02:24 |  | 9.4649 | 20295.249 | FALSE | 48.6161 |  | 48.6161 |  |  | FALSE | 11.0015 | 1139502.7817 | 16.4602 | FALSE
| Amide Hep120220 T0A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 24 | 7010_1_021021023.D | Sample |  | 2021-02-11 02:58 |  | 9.4649 | 43724.8136 | FALSE | 81.5675 |  | 81.5675 |  | 6.8074 | FALSE | 11.0015 | 1467241.8316 | 16.4948 | FALSE
| Amide Hep120220 T0A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 25 | 7010_1_021021024.D | Sample |  | 2021-02-11 03:31 |  | 9.4649 | 108958.1932 | TRUE | 320.4129 |  | 320.4129 |  | 4.1995 | FALSE | 11.0015 | 949650.797 | 15.7524 | FALSE
| Amide Hep120220 T0A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 26 | 7010_1_021021025.D | Sample |  | 2021-02-11 04:04 |  | 9.4649 | 7461.2622 | FALSE | 20.6365 |  | 20.6365 |  |  | FALSE | 11.0057 | 984620.2053 | 16.5316 | FALSE
| Amide Hep120220 T0A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 27 | 7010_1_021021026.D | Sample |  | 2021-02-11 04:37 |  | 9.4649 | 36528.9333 | FALSE | 99.5326 |  | 99.5326 |  |  | FALSE | 11.0015 | 1006033.0484 | 16.7534 | FALSE
| Amide Hep120220 T0A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 28 | 7010_1_021021027.D | Sample |  | 2021-02-11 05:11 |  | 9.4649 | 175912.0604 | TRUE | 325.0104 |  | 325.0104 |  | 2.8616 | FALSE | 11.0015 | 1512105.363 | 16.9221 | FALSE
| Amide Hep120220 T0A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 29 | 7010_1_021021028.D | Sample |  | 2021-02-11 05:44 |  | 9.4649 | 107186.677 | FALSE | 252.5863 |  | 252.5863 |  | 5.3093 | FALSE | 11.0015 | 1178283.8521 | 16.2921 | FALSE
| Amide Hep120220 TMidA1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 30 | 7010_1_021021029.D | Sample |  | 2021-02-11 06:17 |  | 9.4577 | 10060.3721 | FALSE | 22.5564 |  | 22.5564 |  |  | FALSE | 11.0015 | 1214804.9822 | 16.4161 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021030.D | QC | 2 | 2021-02-11 06:50 | 4.4 | 9.4577 | 2433.8256 | FALSE | 4.0576 |  | 4.0576 | 92.2189 | 27.4879 | FALSE | 11.0015 | 1631225.3389 | 16.7198 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021031.D | QC | 11 | 2021-02-11 07:24 | 305 | 9.4577 | 81739.0424 | FALSE | 188.6516 |  | 188.6516 | 61.853 | 3.5925 | FALSE | 11.0015 | 1196600.3583 | 16.9051 | FALSE
| Amide Hep120220 TMidA2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 31 | 7010_1_021021032.D | Sample |  | 2021-02-11 07:57 |  | 9.4649 | 23569.6179 | FALSE | 78.2997 |  | 78.2997 |  |  | FALSE | 11.0015 | 823692.1843 | 16.6719 | FALSE
| Amide Hep120220 TMidA3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 32 | 7010_1_021021033.D | Sample |  | 2021-02-11 08:30 |  | 9.4649 | 7917.2825 | FALSE | 22.4281 |  | 22.4281 |  |  | FALSE | 11.0015 | 961483.9696 | 16.0318 | FALSE
| Amide Hep120220 TMidA4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 33 | 7010_1_021021034.D | Sample |  | 2021-02-11 09:03 |  | 9.4649 | 13793.2742 | FALSE | 28.8761 |  | 28.8761 |  |  | FALSE | 10.9972 | 1301720.7842 | 16.4131 | FALSE
| Amide Hep120220 TMidA5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 34 | 7010_1_021021035.D | Sample |  | 2021-02-11 09:37 |  | 9.4649 | 14485.9955 | FALSE | 28.896 |  | 28.896 |  |  | FALSE | 11.0015 | 1366155.6963 | 16.6035 | FALSE
| Amide Hep120220 TMidA6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 35 | 7010_1_021021036.D | Sample |  | 2021-02-11 10:10 |  | 9.4649 | 9108.6803 | FALSE | 18.6798 |  | 18.6798 |  |  | FALSE | 11.0015 | 1327721.2641 | 16.5042 | FALSE
| Amide Hep120220 TMidA7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 36 | 7010_1_021021037.D | Sample |  | 2021-02-11 10:43 |  | 9.4577 | 6671.8807 | FALSE | 10.8969 |  | 10.8969 |  |  | FALSE | 10.9972 | 1666051.7762 | 16.5629 | FALSE
| Amide Hep120220 TMidA8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 37 | 7010_1_021021038.D | Sample |  | 2021-02-11 11:16 |  | 9.4649 | 4675.233 | TRUE | 7.7972 |  | 7.7972 |  |  | FALSE | 10.9972 | 1631164.8763 | 16.5653 | FALSE
| Amide Hep120220 TMidA9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 38 | 7010_1_021021039.D | Sample |  | 2021-02-11 11:50 |  | 9.4577 | 11011.373 | FALSE | 23.1666 |  | 23.1666 |  |  | FALSE | 11.0015 | 1294679.5364 | 16.1028 | FALSE
| Amide Hep120220 TMidA10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 39 | 7010_1_021021040.D | Sample |  | 2021-02-11 12:23 |  | 9.4649 | 4596.54 | FALSE | 8.1927 |  | 8.1927 |  |  | FALSE | 10.9972 | 1526331.4613 | 16.5777 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021041.D | QC | 7 | 2021-02-11 12:56 | 47 | 9.4577 | 31518.8362 | TRUE | 53.8968 |  | 53.8968 | 114.6741 |  | FALSE | 11.0015 | 1596975.2223 | 16.2858 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021021042.D | QC | 14 | 2021-02-11 13:29 | 1250 | 9.4577 | 308013.0396 | TRUE | 683.1757 |  | 683.1757 | 54.6541 | 4.6124 | FALSE | 10.9972 | 1299106.2208 | 15.9401 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021043.D | QC | 15 | 2021-02-11 14:03 | 2000 | 9.4577 | 1121103.2898 | TRUE | 2268.7618 |  | 2268.7618 | 113.4381 | 4.692 | FALSE | 10.9972 | 1653672.9933 | 16.2989 | FALSE
| Amide Hep120220 TMidA11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 40 | 7010_1_021021044.D | Sample |  | 2021-02-11 14:36 |  | 9.4506 | 6013.1705 | FALSE | 11.6953 |  | 11.6953 |  |  | FALSE | 10.9972 | 1399153.6007 | 16.5182 | FALSE
| Amide Hep120220 TMidA12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 41 | 7010_1_021021045.D | Sample |  | 2021-02-11 15:09 |  | 9.4577 | 3343.9616 | TRUE | 6.5259 |  | 6.5259 |  |  | FALSE | 10.9972 | 1393815.7663 | 16.3068 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021021046.D | Sample |  | 2021-02-11 15:42 |  | 9.4649 | 950.3546 | TRUE | 2.1624 |  | 2.1624 |  |  | FALSE | 10.9972 | 1195003.7494 | 16.3011 | FALSE
| Amide Hep120220 T0B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 51 | 7010_1_021021057.D | Sample |  | 2021-02-11 16:20 |  | 9.4578 | 148693.5364 | TRUE | 390.1618 |  | 390.1618 |  | 5.0955 | FALSE | 10.9972 | 1070637.3302 | 16.1526 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021058.D | Sample |  | 2021-02-11 16:53 |  | 9.4578 | 268481.9312 | TRUE | 651.9171 |  | 651.9171 |  | 7.2184 | FALSE | 10.9972 | 1183429.5555 | 16.487 | FALSE
| Amide Hep120220 T0B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 53 | 7010_1_021021059.D | Sample |  | 2021-02-11 17:26 |  | 9.4577 | 198861.885 | TRUE | 499.488 |  | 499.488 |  | 6.6216 | FALSE | 10.9972 | 1129010.3124 | 16.5765 | FALSE
| Amide Hep120220 T0B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 54 | 7010_1_021021060.D | Sample |  | 2021-02-11 17:59 |  | 9.4577 | 181203.9172 | TRUE | 416.9981 |  | 416.9981 |  | 5.1172 | FALSE | 10.9972 | 1223561.0281 | 16.7504 | FALSE
| Amide Hep120220 T0B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 55 | 7010_1_021021061.D | Sample |  | 2021-02-11 18:32 |  | 9.4577 | 169795.4018 | TRUE | 427.697 |  | 427.697 |  | 4.5618 | FALSE | 10.9972 | 1118870.834 | 16.1968 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021062.D | Sample |  | 2021-02-11 19:06 |  | 9.4577 | 195280.6023 | TRUE | 380.0191 |  | 380.0191 |  | 5.5129 | FALSE | 10.9972 | 1442355.8709 | 16.7947 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021063.D | Sample |  | 2021-02-11 19:39 |  | 9.4577 | 139175.4935 | TRUE | 361.5477 |  | 361.5477 |  | 5.242 | FALSE | 10.9972 | 1078776.6744 | 16.1556 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021064.D | QC | 3 | 2021-02-11 20:12 | 7.1 | 9.4577 | 4202.2807 | TRUE | 7.5937 |  | 7.5937 | 106.9529 | 1.7202 | FALSE | 10.9972 | 1505421.7831 | 16.0836 | FALSE
| Amide Hep120220 T240B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 72 | 7010_1_021021065.D | Sample |  | 2021-02-11 20:49 |  | 9.3438 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9972 | 1368939.113 | 16.2463 | FALSE
| Amide Hep120220 T240B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 73 | 7010_1_021021066.D | Sample |  | 2021-02-11 21:22 |  | 9.4506 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9972 | 1455996.9248 | 16.771 | FALSE
| Amide Hep120220 T240B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 74 | 7010_1_021021067.D | Sample |  | 2021-02-11 21:55 |  | 9.4577 | 10587.4091 | TRUE | 21.6115 |  | 21.6115 |  | 6.3067 | FALSE | 10.993 | 1334235.7534 | 16.2788 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021068.D | Sample |  | 2021-02-11 22:28 |  | 9.4577 | 16603.9999 | FALSE | 36.488 |  | 36.488 |  | 3.8412 | FALSE | 10.9972 | 1240868.7448 | 15.9939 | FALSE
| Amide Hep120220 T240B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 76 | 7010_1_021021069.D | Sample |  | 2021-02-11 23:01 |  | 9.4578 | 54097.1004 | TRUE | 123.7554 |  | 123.7554 |  | 5.5311 | FALSE | 10.993 | 1200683.2968 | 15.7968 | FALSE
| Amide Hep120220 T240B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 77 | 7010_1_021021070.D | Sample |  | 2021-02-11 23:35 |  | 9.4577 | 59449.4129 | TRUE | 154.8157 |  | 154.8157 |  | 4.8406 | FALSE | 10.993 | 1057498.623 | 16.4488 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021071.D | Sample |  | 2021-02-12 00:08 |  | 9.4578 | 255577.3421 | TRUE | 613.4642 |  | 613.4642 |  | 4.9753 | FALSE | 10.9972 | 1193151.4602 | 16.1777 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021072.D | QC | 6 | 2021-02-12 00:41 | 29 | 9.4577 | 17998.2064 | FALSE | 34.7751 |  | 34.7751 | 119.9142 | 21.877 | FALSE | 10.993 | 1411114.9478 | 16.5335 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021073.D | QC | 12 | 2021-02-12 01:14 | 488 | 9.4577 | 225683.5996 | TRUE | 438.6632 |  | 438.6632 | 89.89 | 4.9006 | FALSE | 10.993 | 1451334.565 | 16.7815 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021021074.D | Sample |  | 2021-02-12 01:48 |  | 9.4577 | 203361.9927 | TRUE | 364.9945 |  | 364.9945 |  | 4.9986 | FALSE | 10.9972 | 1561871.923 | 16.8299 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021021075.D | Sample |  | 2021-02-12 02:21 |  | 9.4577 | 169992.9759 | TRUE | 386.3024 |  | 386.3024 |  | 4.6739 | FALSE | 10.993 | 1235820.4763 | 17.052 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021021076.D | Sample |  | 2021-02-12 02:54 |  | 9.4577 | 132557.0747 | TRUE | 296.5448 |  | 296.5448 |  | 4.5472 | FALSE | 10.9972 | 1245796.3823 | 17.1506 | FALSE
| Amide Hep120220 T240B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 78 | 7010_1_021021077.D | Sample |  | 2021-02-12 03:27 |  | 9.4577 | 53141.1413 | FALSE | 132.756 |  | 132.756 |  | 3.512 | FALSE | 10.993 | 1100328.3884 | 16.9299 | FALSE
| Amide Hep120220 T240B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 79 | 7010_1_021021078.D | Sample |  | 2021-02-12 04:01 |  | 9.4578 | 53581.4445 | FALSE | 122.6139 |  | 122.6139 |  | 3.3666 | FALSE | 10.993 | 1200195.6816 | 17.4207 | FALSE
| Amide Hep120220 T240B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 80 | 7010_1_021021079.D | Sample |  | 2021-02-12 04:34 |  | 9.4577 | 43663.0005 | FALSE | 103.7794 |  | 103.7794 |  | 4.3756 | FALSE | 10.993 | 1153710.9073 | 17.2878 | FALSE
| Amide Hep120220 T0C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 81 | 7010_1_021021080.D | Sample |  | 2021-02-12 05:07 |  | 9.4578 | 33540.698 | TRUE | 91.7068 |  | 91.7068 |  | 4.8445 | FALSE | 10.993 | 1001908.3613 | 16.7672 | FALSE
| Amide Hep120220 T0C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 82 | 7010_1_021021081.D | Sample |  | 2021-02-12 05:40 |  | 9.4577 | 20933.8565 | TRUE | 58.8096 |  | 58.8096 |  | 3.742 | FALSE | 10.993 | 972455.3572 | 16.7866 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021082.D | QC | 7 | 2021-02-12 06:14 | 47 | 9.4577 | 18396.0861 | TRUE | 46.3963 |  | 46.3963 | 98.7155 | 19.3118 | FALSE | 10.993 | 1082089.2888 | 16.6215 | FALSE
| Amide Hep120220 T0C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 83 | 7010_1_021021083.D | Sample |  | 2021-02-12 06:47 |  | 9.4577 | 31125.8238 | TRUE | 102.6971 |  | 102.6971 |  | 5.1945 | FALSE | 10.993 | 831032.5682 | 16.8694 | FALSE
| Amide Hep120220 T0C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 84 | 7010_1_021021084.D | Sample |  | 2021-02-12 07:20 |  | 9.4578 | 43083.0735 | TRUE | 114.4647 |  | 114.4647 |  | 4.9227 | FALSE | 10.993 | 1033039.7387 | 16.4451 | FALSE
| Amide Hep120220 T0C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 85 | 7010_1_021021085.D | Sample |  | 2021-02-12 07:53 |  | 9.4577 | 40397.7645 | TRUE | 122.7201 |  | 122.7201 |  | 4.9368 | FALSE | 10.993 | 904113.176 | 16.4554 | FALSE
| Amide Hep120220 T0C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 86 | 7010_1_021021086.D | Sample |  | 2021-02-12 08:27 |  | 9.4577 | 26853.6962 | TRUE | 78.8649 |  | 78.8649 |  | 5.6174 | FALSE | 10.993 | 931779.9049 | 16.8305 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021087.D | QC | 10 | 2021-02-12 09:00 | 191 | 9.4577 | 75096.0524 | TRUE | 162.9609 |  | 162.9609 | 85.3198 | 3.7663 | FALSE | 10.993 | 1269923.3239 | 16.2772 | FALSE
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021088.D | MatrixBlank |  | 2021-02-12 09:33 |  | 9.415 | 10819.8192 | TRUE |  |  |  |  | 311.0493 | FALSE | 11.0015 | 0 |  | FALSE
| Amide Hep120220 T0C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 87 | 7010_1_021021089.D | Sample |  | 2021-02-12 10:06 |  | 9.4577 | 49977.0924 | TRUE | 151.0735 |  | 151.0735 |  | 4.2341 | FALSE | 10.993 | 910738.7204 | 16.7235 | FALSE
| Amide Hep120220 T0C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 88 | 7010_1_021021090.D | Sample |  | 2021-02-12 10:39 |  | 9.4506 | 4828.2009 | FALSE | 8.658 |  | 8.658 |  | 3.7548 | FALSE | 10.993 | 1517162.4518 | 16.8494 | FALSE
| Amide Hep120220 T0C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 89 | 7010_1_021021091.D | Sample |  | 2021-02-12 11:21 |  | 9.4578 | 5722.7741 | FALSE | 11.5309 |  | 11.5309 |  | 9.8475 | FALSE | 10.993 | 1350538.834 | 16.6547 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021092.D | MatrixBlank |  | 2021-02-12 11:54 |  | 9.4649 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.993 | 1537915.2315 | 16.4698 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021093.D | QC | 15 | 2021-02-12 12:28 | 2000 | 9.4577 | 1027457.4399 | TRUE | 2399.9772 |  | 2399.9772 | 119.9989 | 3.929 | FALSE | 10.993 | 1452077.3799 | 16.2193 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021094.D | QC | 14 | 2021-02-12 13:01 | 1250 | 9.4578 | 610067.2346 | TRUE | 1236.3795 |  | 1236.3795 | 98.9104 | 3.7755 | FALSE | 10.993 | 1494237.8182 | 16.518 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021095.D | QC | 13 | 2021-02-12 13:34 | 781 | 9.4577 | 404752.4041 | TRUE | 957.2095 |  | 957.2095 | 122.562 | 4.2922 | FALSE | 10.9887 | 1248386.5293 | 16.8735 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021096.D | QC | 12 | 2021-02-12 14:07 | 488 | 9.4577 | 254052.4579 | TRUE | 543.323 |  | 543.323 | 111.3367 | 4.8693 | FALSE | 10.993 | 1331012.2455 | 16.383 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021097.D | QC | 11 | 2021-02-12 14:40 | 305 | 9.4578 | 146104.8997 | TRUE | 258.6089 |  | 258.6089 | 84.7898 | 4.2981 | FALSE | 10.993 | 1569499.7597 | 16.8251 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021098.D | QC | 10 | 2021-02-12 15:14 | 191 | 9.4578 | 92056.9785 | TRUE | 156.4565 |  | 156.4565 | 81.9144 | 4.7923 | FALSE | 10.993 | 1620578.5035 | 16.6787 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021099.D | QC | 9 | 2021-02-12 15:47 | 119 | 9.4577 | 62918.0767 | TRUE | 123.8681 |  | 123.8681 | 104.0908 | 5.6073 | FALSE | 10.993 | 1395207.6913 | 16.5331 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021100.D | QC | 8 | 2021-02-12 16:20 | 75 | 9.4577 | 37100.8882 | TRUE | 70.8531 |  | 70.8531 | 94.4708 | 5.3287 | FALSE | 10.993 | 1431953.6932 | 16.6109 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021101.D | QC | 7 | 2021-02-12 16:53 | 47 | 9.4577 | 30038.1428 | FALSE | 47.4472 |  | 47.4472 | 100.9516 | 5.6644 | FALSE | 10.993 | 1727908.7341 | 16.5044 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021102.D | QC | 6 | 2021-02-12 17:27 | 29 | 9.4577 | 21197.3761 | TRUE | 33.4452 |  | 33.4452 | 115.3282 | 4.5753 | TRUE | 10.993 | 1727836.1962 | 16.412 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021103.D | QC | 5 | 2021-02-12 18:00 | 18 | 9.4506 | 11111.5883 | TRUE | 18.8528 |  | 18.8528 | 104.7376 | 8.2932 | FALSE | 10.993 | 1604833.3904 | 16.5686 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021104.D | QC | 4 | 2021-02-12 18:33 | 11 | 9.4577 | 8203.8765 | TRUE | 14.7845 |  | 14.7845 | 134.4047 | 1.816 | FALSE | 10.993 | 1510408.5477 | 16.4124 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021105.D | QC | 3 | 2021-02-12 19:07 | 7.1 | 9.4506 | 4949.6632 | FALSE | 9.44 |  | 9.44 | 132.9576 | 7.39 | FALSE | 10.9887 | 1426574.3777 | 16.5389 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021106.D | QC | 2 | 2021-02-12 19:40 | 4.4 | 9.4506 | 3626.9691 | FALSE | 5.9398 |  | 5.9398 | 134.9949 | 30.8576 | FALSE | 10.993 | 1660880.0813 | 16.5037 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021107.D | QC | 1 | 2021-02-12 20:13 | 2.8 | 9.4506 | 1381.743 | TRUE | 2.4808 |  | 2.4808 | 88.5996 |  | FALSE | 10.9887 | 1514529.3891 | 16.5007 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021108.D | QC | 3 | 2021-02-12 20:46 | 7.1 | 9.4577 | 2098.4598 | TRUE | 3.1488 |  | 3.1488 | 44.3486 | 21.5973 | FALSE | 10.9887 | 1812284.3058 | 16.6834 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021109.D | MatrixBlank |  | 2021-02-12 21:20 |  | 9.5005 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9887 | 1876327.1648 | 16.3154 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021110.D | Sample |  | 2021-02-12 21:53 |  | 9.4506 | 15983.9908 | FALSE | 33.6097 |  | 33.6097 |  | 5.672 | FALSE | 10.9887 | 1296524.3149 | 16.4413 | FALSE
| Amide Hep120220 T240C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 102 | 7010_1_021021111.D | Sample |  | 2021-02-12 22:26 |  | 9.4578 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.993 | 1061413.605 | 16.4233 | FALSE
| Amide Hep120220 T240C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 103 | 7010_1_021021112.D | Sample |  | 2021-02-12 22:59 |  | 9.4506 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.993 | 2899512.1558 | 17.9619 | FALSE
| Amide Hep120220 T240C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 104 | 7010_1_021021113.D | Sample |  | 2021-02-12 23:33 |  | 9.4506 | 12833.0242 | TRUE | 11.6476 |  | 11.6476 |  | 9.8762 | FALSE | 10.993 | 2998218.9334 | 17.8928 | FALSE
| Amide Hep120220 T240C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 105 | 7010_1_021021114.D | Sample |  | 2021-02-13 00:06 |  | 9.4506 | 17686.1809 | FALSE | 20.2255 |  | 20.2255 |  | 9.014 | FALSE | 10.9887 | 2381288.7606 | 17.9741 | FALSE
| Amide Hep120220 T240C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 106 | 7010_1_021021115.D | Sample |  | 2021-02-13 00:39 |  | 9.4506 | 19030.2802 | TRUE | 38.7679 |  | 38.7679 |  | 5.8228 | FALSE | 10.9887 | 1338808.5188 | 16.8042 | FALSE
| Amide Hep120220 T240C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 107 | 7010_1_021021116.D | Sample |  | 2021-02-13 01:12 |  | 9.4506 | 25929.1117 | TRUE | 49.2649 |  | 49.2649 |  | 5.8604 | FALSE | 10.9887 | 1436728.2402 | 16.4115 | FALSE
| Amide Hep120220 T240C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 108 | 7010_1_021021117.D | Sample |  | 2021-02-13 01:46 |  | 9.4506 | 24638.7487 | TRUE | 46.3972 |  | 46.3972 |  | 5.1177 | FALSE | 10.9887 | 1449263.6916 | 16.4455 | FALSE
| Amide Hep120220 T240C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 109 | 7010_1_021021118.D | Sample |  | 2021-02-13 02:19 |  | 9.4506 | 26467.189 | TRUE | 59.6947 |  | 59.6947 |  | 7.3168 | FALSE | 10.9887 | 1211357.7857 | 16.428 | FALSE
| Amide Hep120220 T240C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 110 | 7010_1_021021119.D | Sample |  | 2021-02-13 02:52 |  | 9.4506 | 18048.9463 | TRUE | 50.5368 |  | 50.5368 |  | 6.7665 | FALSE | 10.993 | 975022.2483 | 15.7866 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021120.D | QC | 11 | 2021-02-13 03:25 | 305 | 9.4578 | 88205.9928 | TRUE | 123.1622 |  | 123.1622 | 40.3811 | 5.8683 | FALSE | 10.9887 | 1967061.2052 | 16.5843 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021521001.D | QC | 11 | 2021-02-15 09:38 | 305 | 9.4578 | 160794.7256 | TRUE | 415.427 |  | 415.427 | 136.2056 | 5.85 | FALSE | 10.9888 | 1089709.628 | 15.8899 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521002.D | QC | 15 | 2021-02-15 10:11 | 2000 | 9.4579 | 1292278.7706 | TRUE | 2548.9987 |  | 2548.9987 | 127.4499 | 5.2508 | FALSE | 10.9888 | 1746420.9306 | 16.7178 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521003.D | QC | 15 | 2021-02-15 10:44 | 2000 | 9.4507 | 1050232.3441 | TRUE | 1774.3048 |  | 1774.3048 | 88.7152 | 4.7876 | FALSE | 10.9888 | 1885916.7109 | 16.4804 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021521004.D | Sample |  | 2021-02-15 11:17 |  | 9.4507 | 263359.3744 | TRUE | 524.073 |  | 524.073 |  | 6.445 | FALSE | 10.9888 | 1428073.1821 | 16.0304 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021521005.D | Sample |  | 2021-02-15 11:50 |  | 9.4507 | 213673.1861 | TRUE | 468.2724 |  | 468.2724 |  | 5.2063 | FALSE | 10.9888 | 1290491.2495 | 15.8517 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021521006.D | Sample |  | 2021-02-15 12:24 |  | 9.4507 | 163830.7499 | TRUE | 384.3426 |  | 384.3426 |  | 5.5983 | FALSE | 10.9888 | 1196894.7839 | 16.3618 | FALSE
| Amide Hep120220 TMidB1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 60 | 7010_1_021521007.D | Sample |  | 2021-02-15 12:57 |  | 9.4507 | 104942.1315 | TRUE | 214.8822 |  | 214.8822 |  | 5.4481 | FALSE | 10.9888 | 1351722.6553 | 16.3029 | FALSE
| Amide Hep120220 TMidB2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 61 | 7010_1_021521008.D | Sample |  | 2021-02-15 13:30 |  | 9.4579 | 68246.7887 | TRUE | 126.475 |  | 126.475 |  | 4.9471 | FALSE | 10.9888 | 1482501.358 | 15.7608 | FALSE
| Amide Hep120220 TMidB3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 62 | 7010_1_021521009.D | Sample |  | 2021-02-15 14:03 |  | 9.4507 | 110708.0447 | TRUE | 193.144 |  | 193.144 |  | 5.1256 | FALSE | 10.9889 | 1583586.7699 | 16.5115 | FALSE
| Amide Hep120220 TMidB4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 63 | 7010_1_021521010.D | Sample |  | 2021-02-15 14:37 |  | 9.4507 | 95738.1271 | TRUE | 192.2322 |  | 192.2322 |  | 5.1773 | FALSE | 10.9889 | 1375845.0535 | 15.8883 | FALSE
| Amide Hep120220 TMidB5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 64 | 7010_1_021521011.D | Sample |  | 2021-02-15 15:10 |  | 9.4507 | 83366.2738 | TRUE | 153.8725 |  | 153.8725 |  | 5.5962 | FALSE | 10.9888 | 1491909.1765 | 16.5452 | FALSE
| Amide Hep120220 TMidB6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 65 | 7010_1_021521012.D | Sample |  | 2021-02-15 15:43 |  | 9.4507 | 96441.5319 | TRUE | 192.3166 |  | 192.3166 |  | 4.8852 | FALSE | 10.9888 | 1385355.4716 | 16.0732 | FALSE
| Amide Hep120220 TMidB7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 66 | 7010_1_021521013.D | Sample |  | 2021-02-15 16:17 |  | 9.4507 | 41010.0941 | TRUE | 74.6949 |  | 74.6949 |  | 5.3051 | FALSE | 10.9888 | 1501903.8251 | 15.9211 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021521014.D | QC | 4 | 2021-02-15 16:50 | 11 | 9.4507 | 6103.5775 | FALSE | 10.3937 |  | 10.3937 | 94.4881 | 20.7914 | FALSE | 10.9846 | 1597862.8245 | 16.0039 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521015.D | MatrixBlank |  | 2021-02-15 17:23 |  | 9.4223 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9846 | 2065444.3364 | 16.3871 | FALSE
| Amide Hep120220 TMidB8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 67 | 7010_1_021521016.D | Sample |  | 2021-02-15 17:56 |  | 9.4507 | 42784.3638 | TRUE | 76.5031 |  | 76.5031 |  | 6.9487 | FALSE | 10.9846 | 1530077.5364 | 16.7428 | FALSE
| Amide Hep120220 TMidB9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 68 | 7010_1_021521017.D | Sample |  | 2021-02-15 18:30 |  | 9.4507 | 31988.5729 | TRUE | 55.8443 |  | 55.8443 |  | 5.1937 | FALSE | 10.9846 | 1564508.1695 | 16.4263 | FALSE
| Amide Hep120220 TMidB10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 69 | 7010_1_021521018.D | Sample |  | 2021-02-15 19:03 |  | 9.4507 | 13261.0033 | FALSE | 22.6965 |  | 22.6965 |  | 4.2183 | FALSE | 10.9889 | 1591418.6379 | 16.5783 | FALSE
| Amide Hep120220 TMidB11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 70 | 7010_1_021521019.D | Sample |  | 2021-02-15 19:36 |  | 9.4507 | 12246.3439 | FALSE | 21.4336 |  | 21.4336 |  | 6.2031 | FALSE | 10.9888 | 1556081.2329 | 16.1663 | FALSE
| Amide Hep120220 TMidB12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 71 | 7010_1_021521020.D | Sample |  | 2021-02-15 20:09 |  | 9.4507 | 5342.6433 | FALSE | 13.3509 |  | 13.3509 |  |  | FALSE | 10.9846 | 1089125.1205 | 15.782 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021521021.D | QC | 6 | 2021-02-15 20:43 | 29 | 9.4507 | 17354.812 | TRUE | 42.84 |  | 42.84 | 147.724 | 7.7796 | FALSE | 10.9846 | 1105257.3151 | 15.5945 | FALSE
| Amide Hep120220 TMidC1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 90 | 7010_1_021521022.D | Sample |  | 2021-02-15 21:16 |  | 9.4507 | 12485.8446 | TRUE | 25.4321 |  | 25.4321 |  | 5.9688 | FALSE | 10.9846 | 1337523.5952 | 15.7728 | FALSE
| Amide Hep120220 TMidC2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 91 | 7010_1_021521023.D | Sample |  | 2021-02-15 21:49 |  | 9.4507 | 37450.9895 | TRUE | 78.7505 |  | 78.7505 |  | 5.2843 | FALSE | 10.9846 | 1301363.9414 | 15.9555 | FALSE
| Amide Hep120220 TMidC3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 92 | 7010_1_021521024.D | Sample |  | 2021-02-15 22:22 |  | 9.4507 | 9188.2981 | TRUE | 19.9974 |  | 19.9974 |  | 5.8889 | FALSE | 10.9846 | 1251213.9398 | 15.6121 | FALSE
| Amide Hep120220 TMidC4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 93 | 7010_1_021521025.D | Sample |  | 2021-02-15 22:56 |  | 9.4507 | 3744.2492 | TRUE | 10.0853 |  | 10.0853 |  |  | FALSE | 10.9846 | 1010158.4132 | 15.6893 | FALSE
| Amide Hep120220 TMidC5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 94 | 7010_1_021521026.D | Sample |  | 2021-02-15 23:29 |  | 9.4507 | 6052.3225 | TRUE | 16.962 |  | 16.962 |  | 6.6871 | FALSE | 10.9846 | 971417.7653 | 15.9506 | FALSE
| Amide Hep120220 TMidC6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 95 | 7010_1_021521027.D | Sample |  | 2021-02-16 00:02 |  | 9.4507 | 2038.2675 | FALSE | 4.0808 |  | 4.0808 |  | 5.5944 | FALSE | 10.9846 | 1358371.3736 | 16.5259 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021521028.D | QC | 14 | 2021-02-16 00:35 | 1250 | 9.4507 | 273811.563 | TRUE | 661.0574 |  | 661.0574 | 52.8846 | 4.1924 | FALSE | 10.9846 | 1191185.6119 | 16.177 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521029.D | QC | 15 | 2021-02-16 01:09 | 2000 | 9.4507 | 1180063.7186 | FALSE | 3447.0564 |  | 3447.0564 | 172.3528 | 5.2682 | FALSE | 10.9846 | 1301790.9075 | 16.6447 | FALSE
| Amide Hep120220 TMidC7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 96 | 7010_1_021521030.D | Sample |  | 2021-02-16 01:42 |  | 9.4507 | 5660.442 | FALSE | 11.1945 |  | 11.1945 |  |  | FALSE | 10.9337 | 1375933.0656 | 16.6485 | TRUE
| Amide Hep120220 TMidC8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 97 | 7010_1_021521031.D | Sample |  | 2021-02-16 02:15 |  | 9.4507 | 4163.7332 | TRUE | 8.0011 |  | 8.0011 |  | 7.0786 | FALSE | 10.9888 | 1415696.2804 | 16.9998 | FALSE
| Amide Hep120220 TMidC9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 98 | 7010_1_021521032.D | Sample |  | 2021-02-16 02:48 |  | 9.4507 | 4538.2963 | TRUE | 10.4916 |  | 10.4916 |  | 9.6418 | FALSE | 10.9846 | 1177007.9129 | 16.2408 | FALSE
| Amide Hep120220 TMidC10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 99 | 7010_1_021521033.D | Sample |  | 2021-02-16 03:22 |  | 9.4507 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9846 | 1573056.5806 | 16.6627 | FALSE
| Amide Hep120220 TMidC11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 100 | 7010_1_021521034.D | Sample |  | 2021-02-16 03:55 |  | 9.4507 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9846 | 1487167.1693 | 16.3383 | FALSE
| Amide Hep120220 TMidC12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 101 | 7010_1_021521035.D | Sample |  | 2021-02-16 04:28 |  | 9.4579 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9846 | 1221022.527 | 16.1578 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021521036.D | Sample |  | 2021-02-16 05:01 |  | 9.4579 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9846 | 1455972.8968 | 16.8817 | FALSE
| Amide Hep120220 T240A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 43 | 7010_1_021521037.D | Sample |  | 2021-02-16 05:35 |  | 9.4507 | 2398.2458 | TRUE | 5.7196 |  | 5.7196 |  | 10.481 | TRUE | 10.9846 | 1140475.4058 | 16.5723 | FALSE
| Amide Hep120220 T240A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 44 | 7010_1_021521038.D | Sample |  | 2021-02-16 06:08 |  | 9.4507 | 5159.7672 | FALSE | 12.7178 |  | 12.7178 |  | 6.5247 | TRUE | 10.9846 | 1104145.2921 | 16.3082 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021521039.D | QC | 9 | 2021-02-16 06:41 | 119 | 9.4507 | 58124.7145 | TRUE | 178.1028 |  | 178.1028 | 149.6662 | 5.6706 | FALSE | 10.9846 | 900503.6595 | 16.8534 | FALSE
| Amide Hep120220 T240A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 45 | 7010_1_021521040.D | Sample |  | 2021-02-16 07:15 |  | 9.4507 | 14390.4519 | FALSE | 36.8174 |  | 36.8174 |  | 4.2344 | FALSE | 10.9846 | 1065852.6822 | 16.3473 | FALSE
| Amide Hep120220 T240A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 46 | 7010_1_021521041.D | Sample |  | 2021-02-16 07:48 |  | 9.4507 | 18509.4418 | FALSE | 52.8848 |  | 52.8848 |  |  | FALSE | 10.9846 | 955690.2177 | 16.0574 | FALSE
| Amide Hep120220 T240A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 47 | 7010_1_021521042.D | Sample |  | 2021-02-16 08:21 |  | 9.4507 | 22607.7648 | FALSE | 66.3414 |  | 66.3414 |  | 4.3698 | FALSE | 10.9846 | 931565.1466 | 16.0569 | FALSE
| Amide Hep120220 T240A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 48 | 7010_1_021521043.D | Sample |  | 2021-02-16 08:54 |  | 9.4507 | 28034.1872 | FALSE | 50.1226 |  | 50.1226 |  |  | FALSE | 10.9804 | 1526895.0041 | 15.9114 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521044.D | MatrixBlank |  | 2021-02-16 09:27 |  | 9.4294 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.9803 | 1813919.0893 | 16.2148 | FALSE
| Amide Hep120220 T240A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 49 | 7010_1_021521045.D | Sample |  | 2021-02-16 10:01 |  | 9.4507 | 24319.5098 | FALSE | 63.5633 |  | 63.5633 |  | 4.9734 | FALSE | 10.9804 | 1045655.2677 | 16.0693 | FALSE
| Amide Hep120220 T240A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 50 | 7010_1_021521046.D | Sample |  | 2021-02-16 10:34 |  | 9.4436 | 22458.1789 | FALSE | 56.2594 |  | 56.2594 |  |  | FALSE | 10.9804 | 1090324.8465 | 16.5326 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021521047.D | QC | 7 | 2021-02-16 11:07 | 47 | 9.4507 | 21080.0549 | FALSE | 53.876 |  | 53.876 | 114.6298 | 4.7081 | FALSE | 10.9804 | 1068480.751 | 16.333 | FALSE

## 916
Sample | [8 empty] | 916 Method | 916 Results | [6 empty] | Qualifier (119.0 -> 69.0) Results |  | Qualifier (198.0 -> 66.0) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | Ratio | MI | RT | Resp. | Ratio | MI
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021001.D | MatrixBlank |  | 2021-02-10 14:46 |  | 11.9873 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.0142 | 471.2302 | 18.2619 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021002.D | MatrixBlank |  | 2021-02-10 15:19 |  | 12.1148 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.0057 | 1098520.6023 | 16.617 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021003.D | Cal | 15 | 2021-02-10 15:52 | 2000 | 12.125 | 883354.3379 | FALSE | 1922.793 |  | 1922.793 | 96.1396 | 50.0732 | FALSE | 8.4954 | FALSE | 11.0057 | 2021923.1562 | 17.4464 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021004.D | Cal | 14 | 2021-02-10 16:25 | 1250 | 12.1199 | 633300.1913 | FALSE | 1474.0893 |  | 1474.0893 | 117.9271 | 50.1168 | FALSE | 9.1376 | FALSE | 11.0057 | 1806607.3565 | 17.0202 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021005.D | Cal | 13 | 2021-02-10 16:59 | 781 | 12.125 | 403844.2583 | FALSE | 717.5521 |  | 717.5521 | 91.8761 | 48.2963 | FALSE | 10.1003 | FALSE | 11.0057 | 2201383.4929 | 17.6488 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021006.D | Cal | 12 | 2021-02-10 17:32 | 488 | 12.125 | 203428.2484 | FALSE | 376.9139 |  | 376.9139 | 77.2365 | 49.243 | FALSE | 9.6152 | FALSE | 11.0057 | 2046716.1773 | 17.3102 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021007.D | Cal | 11 | 2021-02-10 18:05 | 305 | 12.125 | 136893.6968 | FALSE | 242.7224 |  | 242.7224 | 79.5811 | 48.8566 | FALSE | 9.6618 | FALSE | 11.0057 | 2113378.3928 | 17.4201 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021008.D | Cal | 10 | 2021-02-10 18:38 | 191 | 12.1199 | 87942.5265 | FALSE | 202.7198 |  | 202.7198 | 106.136 | 49.6681 | FALSE | 9.6214 | FALSE | 11.0057 | 1619841.6805 | 16.9256 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021009.D | Cal | 9 | 2021-02-10 19:12 | 119 | 12.125 | 52027.9929 | FALSE | 92.9301 |  | 92.9301 | 78.0925 | 49.44 | FALSE | 9.0256 | FALSE | 11.0057 | 2070468.6065 | 17.1909 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021010.D | Cal | 8 | 2021-02-10 19:45 | 75 | 12.1199 | 34712.93 | FALSE | 71.8029 |  | 71.8029 | 95.7371 | 56.6508 | FALSE | 9.6055 | FALSE | 11.0015 | 1784585.9395 | 16.7087 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021011.D | Cal | 7 | 2021-02-10 20:18 | 47 | 12.1199 | 18661.0809 | FALSE | 55.46 |  | 55.46 | 117.9999 | 54.616 | FALSE | 5.6827 | FALSE | 11.0015 | 1240301.514 | 17.0934 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021012.D | Cal | 6 | 2021-02-10 20:52 | 29 | 12.1199 | 9891.085 | FALSE | 28.7281 |  | 28.7281 | 99.0626 | 44.3804 | FALSE | 4.4657 | FALSE | 11.0057 | 1266186.0052 | 15.9545 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021013.D | Cal | 5 | 2021-02-10 21:25 | 18 | 12.1199 | 6536.9904 | FALSE | 21.4655 |  | 21.4655 | 119.2529 | 62.8065 | FALSE | 8.0879 | FALSE | 11.0057 | 1119241.1397 | 16.5974 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021014.D | Cal | 4 | 2021-02-10 21:58 | 11 | 12.125 | 2686.9931 | FALSE | 8.4563 |  | 8.4563 | 76.8751 | 51.2072 | FALSE | 10.6871 | FALSE | 11.0015 | 1166503.0569 | 16.437 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021015.D | Cal | 3 | 2021-02-10 22:32 | 7.1 | 12.1199 | 2164.4121 | FALSE | 6.5211 |  | 6.5211 | 91.8471 | 28.9253 | FALSE |  | FALSE | 11.0015 | 1218263.1222 | 16.2663 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021016.D | Cal | 2 | 2021-02-10 23:05 | 4.4 | 12.1199 | 1120.2437 | FALSE | 2.7626 |  | 2.7626 | 62.7868 | 130.8315 | FALSE | 18.823 | FALSE | 11.0015 | 1487904.2374 | 16.5453 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021017.D | Cal | 1 | 2021-02-10 23:38 | 2.8 | 12.1352 | 1172.3802 | TRUE | 3.361 |  | 3.361 | 120.0349 | 38.0939 | FALSE | 14.8035 | FALSE | 11.0015 | 1279997.1288 | 16.3886 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021018.D | QC | 3 | 2021-02-11 00:11 | 7.1 | 12.1199 | 1095.2552 | FALSE | 2.8228 |  | 2.8228 | 39.7584 | 107.5294 | FALSE | 32.7818 | FALSE | 11.0015 | 1423685.6191 | 16.5971 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021019.D | MatrixBlank |  | 2021-02-11 00:45 |  | 12.0486 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.0057 | 1295784.3796 | 16.5578 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021030.D | QC | 2 | 2021-02-11 06:50 | 4.4 | 12.1301 | 1871.6193 | FALSE | 4.2106 |  | 4.2106 | 95.6949 | 77.766 | FALSE | 6.6477 | FALSE | 11.0015 | 1631225.3389 | 16.7198 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021031.D | QC | 11 | 2021-02-11 07:24 | 305 | 12.1199 | 35093.0078 | FALSE | 108.6061 |  | 108.6061 | 35.6085 | 56.6015 | FALSE | 4.11 | FALSE | 11.0015 | 1196600.3583 | 16.9051 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021041.D | QC | 7 | 2021-02-11 12:56 | 47 | 12.1148 | 19364.7002 | FALSE | 44.6555 |  | 44.6555 | 95.0116 | 52.5275 | FALSE | 6.27 | FALSE | 11.0015 | 1596975.2223 | 16.2858 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021021042.D | QC | 14 | 2021-02-11 13:29 | 1250 | 12.1148 | 175442.7914 | FALSE | 518.7121 |  | 518.7121 | 41.497 | 53.5647 | FALSE | 6.9089 | FALSE | 10.9972 | 1299106.2208 | 15.9401 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021043.D | QC | 15 | 2021-02-11 14:03 | 2000 | 12.1148 | 677016.2177 | FALSE | 1774.4845 |  | 1774.4845 | 88.7242 | 54.8273 | FALSE | 5.379 | FALSE | 10.9972 | 1653672.9933 | 16.2989 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021064.D | QC | 3 | 2021-02-11 20:12 | 7.1 | 12.1148 | 2589.92 | FALSE | 6.3146 |  | 6.3146 | 88.9379 |  | FALSE | 6.261 | FALSE | 10.9972 | 1505421.7831 | 16.0836 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021072.D | QC | 6 | 2021-02-12 00:41 | 29 | 12.1148 | 12442.1997 | FALSE | 32.4366 |  | 32.4366 | 111.8504 | 52.941 | FALSE | 5.2259 | FALSE | 10.993 | 1411114.9478 | 16.5335 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021073.D | QC | 12 | 2021-02-12 01:14 | 488 | 12.1097 | 167505.9202 | FALSE | 440.1663 |  | 440.1663 | 90.198 | 51.7391 | FALSE | 7.2258 | FALSE | 10.993 | 1451334.565 | 16.7815 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021087.D | QC | 10 | 2021-02-12 09:00 | 191 | 12.1097 | 52422.0329 | FALSE | 153.4707 |  | 153.4707 | 80.3511 | 51.9593 | FALSE | 5.0934 | FALSE | 10.993 | 1269923.3239 | 16.2772 | FALSE
| ACN | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 1 | 7010_1_021021088.D | MatrixBlank |  | 2021-02-12 09:33 |  | 12.1505 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.0015 | 0 |  | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021082.D | QC | 7 | 2021-02-12 06:14 | 47 | 12.1148 | 14168.668 | FALSE | 48.2351 |  | 48.2351 | 102.6278 | 50.3564 | FALSE | 2.1868 | FALSE | 10.993 | 1082089.2888 | 16.6215 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021092.D | MatrixBlank |  | 2021-02-12 11:54 |  | 12.0944 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 10.993 | 1537915.2315 | 16.4698 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021021093.D | QC | 15 | 2021-02-12 12:28 | 2000 | 12.1097 | 814740.7404 | FALSE | 2680.3034 |  | 2680.3034 | 134.0152 | 52.6123 | FALSE | 6.4562 | FALSE | 10.993 | 1452077.3799 | 16.2193 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 4 | 7010_1_021021094.D | QC | 14 | 2021-02-12 13:01 | 1250 | 12.1097 | 462643.6839 | FALSE | 1277.003 |  | 1277.003 | 102.1602 | 53.6775 | FALSE | 5.7907 | FALSE | 10.993 | 1494237.8182 | 16.518 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 5 | 7010_1_021021095.D | QC | 13 | 2021-02-12 13:34 | 781 | 12.1097 | 256013.6793 | FALSE | 808.9653 |  | 808.9653 | 103.5807 | 55.5258 | FALSE | 5.1319 | FALSE | 10.9887 | 1248386.5293 | 16.8735 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 6 | 7010_1_021021096.D | QC | 12 | 2021-02-12 14:07 | 488 | 12.1097 | 154927.6117 | FALSE | 444.0723 |  | 444.0723 | 90.9984 | 53.4639 | FALSE | 5.8305 | FALSE | 10.993 | 1331012.2455 | 16.383 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021021097.D | QC | 11 | 2021-02-12 14:40 | 305 | 12.1097 | 116529.3716 | FALSE | 279.1114 |  | 279.1114 | 91.5119 | 51.8325 | FALSE | 6.423 | FALSE | 10.993 | 1569499.7597 | 16.8251 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 8 | 7010_1_021021098.D | QC | 10 | 2021-02-12 15:14 | 191 | 12.1097 | 66791.2152 | FALSE | 153.2247 |  | 153.2247 | 80.2224 | 53.0547 | FALSE | 7.2043 | FALSE | 10.993 | 1620578.5035 | 16.6787 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021021099.D | QC | 9 | 2021-02-12 15:47 | 119 | 12.1097 | 39384.0739 | FALSE | 104.4981 |  | 104.4981 | 87.8136 | 54.103 | FALSE | 7.1605 | FALSE | 10.993 | 1395207.6913 | 16.5331 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 10 | 7010_1_021021100.D | QC | 8 | 2021-02-12 16:20 | 75 | 12.1097 | 30785.3734 | FALSE | 79.413 |  | 79.413 | 105.8839 | 54.0409 | FALSE | 6.777 | FALSE | 10.993 | 1431953.6932 | 16.6109 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021021101.D | QC | 7 | 2021-02-12 16:53 | 47 | 12.1097 | 18654.5118 | FALSE | 39.7411 |  | 39.7411 | 84.5555 | 52.8859 | FALSE | 6.7982 | FALSE | 10.993 | 1727908.7341 | 16.5044 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021021102.D | QC | 6 | 2021-02-12 17:27 | 29 | 12.1097 | 11404.6424 | FALSE | 24.2645 |  | 24.2645 | 83.6708 | 53.9385 | FALSE | 8.4826 | FALSE | 10.993 | 1727836.1962 | 16.412 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 13 | 7010_1_021021103.D | QC | 5 | 2021-02-12 18:00 | 18 | 12.1097 | 8211.4302 | FALSE | 18.8008 |  | 18.8008 | 104.4488 | 56.0683 | FALSE | 7.0884 | FALSE | 10.993 | 1604833.3904 | 16.5686 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021021104.D | QC | 4 | 2021-02-12 18:33 | 11 | 12.1097 | 3700.0056 | FALSE | 8.9934 |  | 8.9934 | 81.7585 | 30.9869 | FALSE | 9.4692 | FALSE | 10.993 | 1510408.5477 | 16.4124 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 15 | 7010_1_021021105.D | QC | 3 | 2021-02-12 19:07 | 7.1 | 12.1046 | 1968.3538 | FALSE | 5.0638 |  | 5.0638 | 71.3215 | 38.7487 | FALSE | 15.0796 | FALSE | 10.9887 | 1426574.3777 | 16.5389 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 16 | 7010_1_021021106.D | QC | 2 | 2021-02-12 19:40 | 4.4 | 12.1097 | 1419.5034 | FALSE | 3.1361 |  | 3.1361 | 71.2759 | 136.0172 | FALSE | 18.9486 | FALSE | 10.993 | 1660880.0813 | 16.5037 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 17 | 7010_1_021021107.D | QC | 1 | 2021-02-12 20:13 | 2.8 | 12.1097 | 864.1081 | FALSE | 2.0934 |  | 2.0934 | 74.7637 | 73.4818 | FALSE | 37.6934 | FALSE | 10.9887 | 1514529.3891 | 16.5007 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 18 | 7010_1_021021108.D | QC | 3 | 2021-02-12 20:46 | 7.1 | 12.1097 | 1381.6581 | TRUE | 2.7974 |  | 2.7974 | 39.4004 | 67.2047 | FALSE | 29.4106 | FALSE | 10.9887 | 1812284.3058 | 16.6834 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021021109.D | MatrixBlank |  | 2021-02-12 21:20 |  | 12.2882 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 10.9887 | 1876327.1648 | 16.3154 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 19 | 7010_1_021021120.D | QC | 11 | 2021-02-13 03:25 | 305 | 12.1097 | 58914.012 | FALSE | 110.9358 |  | 110.9358 | 36.3724 | 52.6408 | FALSE | 9.4283 | FALSE | 10.9887 | 1967061.2052 | 16.5843 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 7 | 7010_1_021521001.D | QC | 11 | 2021-02-15 09:38 | 305 | 12.1098 | 94716.1353 | FALSE | 328.1796 |  | 328.1796 | 107.5999 | 55.0046 | FALSE | 5.7237 | FALSE | 10.9888 | 1089709.628 | 15.8899 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521002.D | QC | 15 | 2021-02-15 10:11 | 2000 | 12.1098 | 895881.1983 | FALSE | 2366.864 |  | 2366.864 | 118.3432 | 51.1892 | FALSE | 8.5777 | FALSE | 10.9888 | 1746420.9306 | 16.7178 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521003.D | QC | 15 | 2021-02-15 10:44 | 2000 | 12.1047 | 798526.4181 | FALSE | 1849.399 |  | 1849.399 | 92.47 | 53.2743 | FALSE | 6.1077 | FALSE | 10.9888 | 1885916.7109 | 16.4804 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 14 | 7010_1_021521014.D | QC | 4 | 2021-02-15 16:50 | 11 | 12.1047 | 5476.7147 | FALSE | 12.5873 |  | 12.5873 | 114.4303 | 44.0563 | FALSE | 5.5289 | FALSE | 10.9846 | 1597862.8245 | 16.0039 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521015.D | MatrixBlank |  | 2021-02-15 17:23 |  | 12.1302 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 10.9846 | 2065444.3364 | 16.3871 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 12 | 7010_1_021521021.D | QC | 6 | 2021-02-15 20:43 | 29 | 12.1047 | 9368.0921 | FALSE | 31.1775 |  | 31.1775 | 107.5086 | 56.643 | FALSE | 9.9959 | FALSE | 10.9846 | 1105257.3151 | 15.5945 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\021021 | HC Amides 12/14 | 20 | 7010_1_021521028.D | QC | 14 | 2021-02-16 00:35 | 1250 | 12.1047 | 141377.4599 | FALSE | 453.1726 |  | 453.1726 | 36.2538 | 54.6649 | FALSE | 5.7678 | FALSE | 10.9846 | 1191185.6119 | 16.177 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 3 | 7010_1_021521029.D | QC | 15 | 2021-02-16 01:09 | 2000 | 12.1047 | 695461.9862 | FALSE | 2501.6432 |  | 2501.6432 | 125.0822 | 53.4267 | FALSE | 7.0326 | FALSE | 10.9846 | 1301790.9075 | 16.6447 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 9 | 7010_1_021521039.D | QC | 9 | 2021-02-16 06:41 | 119 | 12.0996 | 31776.2763 | FALSE | 130.9328 |  | 130.9328 | 110.0276 | 60.235 | FALSE | 4.9443 | FALSE | 10.9846 | 900503.6595 | 16.8534 | FALSE
| Method Spike | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 2 | 7010_1_021521044.D | MatrixBlank |  | 2021-02-16 09:27 |  | 12.1404 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 10.9803 | 1813919.0893 | 16.2148 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\021021 | HC Amides 12/7 | 11 | 7010_1_021521047.D | QC | 7 | 2021-02-16 11:07 | 47 | 12.0996 | 15630.3392 | FALSE | 53.9155 |  | 53.9155 | 114.7137 | 48.4926 | FALSE | 5.1408 | FALSE | 10.9804 | 1068480.751 | 16.333 | FALSE
| Amide Hep120220 T0A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 21 | 7010_1_021021020.D | Sample |  | 2021-02-11 01:18 |  | 12.1199 | 3824.1582 | FALSE | 14.3785 |  | 14.3785 |  | 78.8591 | FALSE | 9.0508 | FALSE | 11.0057 | 976881.6283 | 16.4514 | FALSE
| Amide Hep120220 T0A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 22 | 7010_1_021021021.D | Sample |  | 2021-02-11 01:51 |  | 12.1199 | 2485.0326 | FALSE | 7.6488 |  | 7.6488 |  | 68.8789 | FALSE | 18.5098 | FALSE | 11.0057 | 1192634.9348 | 16.585 | FALSE
| Amide Hep120220 T0A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 23 | 7010_1_021021022.D | Sample |  | 2021-02-11 02:24 |  | 12.1199 | 13959.9669 | FALSE | 45.1178 |  | 45.1178 |  | 63.7895 | FALSE | 8.4233 | FALSE | 11.0015 | 1139502.7817 | 16.4602 | FALSE
| Amide Hep120220 T0A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 24 | 7010_1_021021023.D | Sample |  | 2021-02-11 02:58 |  | 12.1199 | 14087.2866 | FALSE | 35.3294 |  | 35.3294 |  | 59.6071 | FALSE | 9.3236 | FALSE | 11.0015 | 1467241.8316 | 16.4948 | FALSE
| Amide Hep120220 T0A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 25 | 7010_1_021021024.D | Sample |  | 2021-02-11 03:31 |  | 12.1199 | 25393.0968 | FALSE | 98.939 |  | 98.939 |  | 62.1558 | FALSE | 6.1243 | FALSE | 11.0015 | 949650.797 | 15.7524 | FALSE
| Amide Hep120220 T0A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 26 | 7010_1_021021025.D | Sample |  | 2021-02-11 04:04 |  | 12.1199 | 9830.2487 | FALSE | 36.7417 |  | 36.7417 |  | 45.9111 | FALSE | 6.3762 | FALSE | 11.0057 | 984620.2053 | 16.5316 | FALSE
| Amide Hep120220 T0A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 27 | 7010_1_021021026.D | Sample |  | 2021-02-11 04:37 |  | 12.1199 | 8831.73 | FALSE | 32.2945 |  | 32.2945 |  | 63.3213 | FALSE | 7.952 | FALSE | 11.0015 | 1006033.0484 | 16.7534 | FALSE
| Amide Hep120220 T0A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 28 | 7010_1_021021027.D | Sample |  | 2021-02-11 05:11 |  | 12.1199 | 19129.7855 | FALSE | 46.5976 |  | 46.5976 |  | 57.4425 | FALSE | 5.5601 | FALSE | 11.0015 | 1512105.363 | 16.9221 | FALSE
| Amide Hep120220 T0A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 29 | 7010_1_021021028.D | Sample |  | 2021-02-11 05:44 |  | 12.1199 | 20309.4772 | FALSE | 63.5808 |  | 63.5808 |  | 52.9106 | FALSE | 8.8368 | FALSE | 11.0015 | 1178283.8521 | 16.2921 | FALSE
| Amide Hep120220 TMidA1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 30 | 7010_1_021021029.D | Sample |  | 2021-02-11 06:17 |  | 12.1199 | 3863.7753 | FALSE | 11.6795 |  | 11.6795 |  | 81.026 | FALSE | 11.0734 | FALSE | 11.0015 | 1214804.9822 | 16.4161 | FALSE
| Amide Hep120220 TMidA2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 31 | 7010_1_021021032.D | Sample |  | 2021-02-11 07:57 |  | 12.1199 | 13263.108 | FALSE | 59.3743 |  | 59.3743 |  | 48.1539 | FALSE | 5.5306 | FALSE | 11.0015 | 823692.1843 | 16.6719 | FALSE
| Amide Hep120220 TMidA3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 32 | 7010_1_021021033.D | Sample |  | 2021-02-11 08:30 |  | 12.1199 | 2362.1811 | FALSE | 9.0197 |  | 9.0197 |  | 68.5791 | FALSE | 8.0253 | FALSE | 11.0015 | 961483.9696 | 16.0318 | FALSE
| Amide Hep120220 TMidA4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 33 | 7010_1_021021034.D | Sample |  | 2021-02-11 09:03 |  | 12.1148 | 6136.5252 | FALSE | 17.3195 |  | 17.3195 |  | 58.786 | FALSE | 6.2466 | FALSE | 10.9972 | 1301720.7842 | 16.4131 | FALSE
| Amide Hep120220 TMidA5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 34 | 7010_1_021021035.D | Sample |  | 2021-02-11 09:37 |  | 12.1199 | 4599.5323 | FALSE | 12.364 |  | 12.364 |  | 63.152 | FALSE | 8.7058 | FALSE | 11.0015 | 1366155.6963 | 16.6035 | FALSE
| Amide Hep120220 TMidA6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 35 | 7010_1_021021036.D | Sample |  | 2021-02-11 10:10 |  | 12.1148 | 2225.0489 | FALSE | 6.151 |  | 6.151 |  | 308.1366 | FALSE | 4.2592 | FALSE | 11.0015 | 1327721.2641 | 16.5042 | FALSE
| Amide Hep120220 TMidA7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 36 | 7010_1_021021037.D | Sample |  | 2021-02-11 10:43 |  | 12.1148 | 1516.6282 | FALSE | 3.3404 |  | 3.3404 |  | 59.2519 | FALSE | 21.0681 | FALSE | 10.9972 | 1666051.7762 | 16.5629 | FALSE
| Amide Hep120220 TMidA8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 37 | 7010_1_021021038.D | Sample |  | 2021-02-11 11:16 |  | 12.1148 | 1642.0348 | FALSE | 3.6941 |  | 3.6941 |  | 117.7527 | FALSE | 9.5533 | FALSE | 10.9972 | 1631164.8763 | 16.5653 | FALSE
| Amide Hep120220 TMidA9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 38 | 7010_1_021021039.D | Sample |  | 2021-02-11 11:50 |  | 12.1148 | 4941.7324 | FALSE | 14.0192 |  | 14.0192 |  | 76.4677 | FALSE | 5.3066 | FALSE | 11.0015 | 1294679.5364 | 16.1028 | FALSE
| Amide Hep120220 TMidA10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 39 | 7010_1_021021040.D | Sample |  | 2021-02-11 12:23 |  | 12.1199 | 3925.5734 | FALSE | 9.4425 |  | 9.4425 |  | 51.9183 | FALSE | 10.5219 | FALSE | 10.9972 | 1526331.4613 | 16.5777 | FALSE
| Amide Hep120220 TMidA11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 40 | 7010_1_021021044.D | Sample |  | 2021-02-11 14:36 |  | 12.1148 | 5516.8832 | FALSE | 14.4828 |  | 14.4828 |  | 57.2093 | FALSE | 6.6676 | FALSE | 10.9972 | 1399153.6007 | 16.5182 | FALSE
| Amide Hep120220 TMidA12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 41 | 7010_1_021021045.D | Sample |  | 2021-02-11 15:09 |  | 12.1148 | 10177.305 | FALSE | 26.8484 |  | 26.8484 |  | 40.2643 | FALSE | 5.9658 | FALSE | 10.9972 | 1393815.7663 | 16.3068 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021021046.D | Sample |  | 2021-02-11 15:42 |  | 12.1148 | 2139.4734 | FALSE | 6.5715 |  | 6.5715 |  | 55.0713 | FALSE | 3.1783 | FALSE | 10.9972 | 1195003.7494 | 16.3011 | FALSE
| Amide Hep120220 T0B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 51 | 7010_1_021021057.D | Sample |  | 2021-02-11 16:20 |  | 12.1148 | 30443.4959 | FALSE | 105.2709 |  | 105.2709 |  | 53.0513 | FALSE | 5.7147 | FALSE | 10.9972 | 1070637.3302 | 16.1526 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021058.D | Sample |  | 2021-02-11 16:53 |  | 12.1148 | 21665.5172 | FALSE | 67.5544 |  | 67.5544 |  | 60.0784 | FALSE | 7.428 | FALSE | 10.9972 | 1183429.5555 | 16.487 | FALSE
| Amide Hep120220 T0B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 53 | 7010_1_021021059.D | Sample |  | 2021-02-11 17:26 |  | 12.1148 | 57245.24 | FALSE | 189.0995 |  | 189.0995 |  | 54.7118 | FALSE | 7.1141 | FALSE | 10.9972 | 1129010.3124 | 16.5765 | FALSE
| Amide Hep120220 T0B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 54 | 7010_1_021021060.D | Sample |  | 2021-02-11 17:59 |  | 12.1148 | 40591.8979 | FALSE | 123.0109 |  | 123.0109 |  | 56.4991 | FALSE | 7.3255 | FALSE | 10.9972 | 1223561.0281 | 16.7504 | FALSE
| Amide Hep120220 T0B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 55 | 7010_1_021021061.D | Sample |  | 2021-02-11 18:32 |  | 12.1148 | 49937.8217 | FALSE | 166.1196 |  | 166.1196 |  | 53.8603 | FALSE | 7.0534 | FALSE | 10.9972 | 1118870.834 | 16.1968 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021062.D | Sample |  | 2021-02-11 19:06 |  | 12.1148 | 61139.7972 | FALSE | 157.6522 |  | 157.6522 |  | 53.5311 | FALSE | 9.2392 | FALSE | 10.9972 | 1442355.8709 | 16.7947 | FALSE
| Amide Hep120220 T0B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 56 | 7010_1_021021063.D | Sample |  | 2021-02-11 19:39 |  | 12.1148 | 47119.6901 | FALSE | 162.5193 |  | 162.5193 |  | 57.5867 | FALSE | 5.6038 | FALSE | 10.9972 | 1078776.6744 | 16.1556 | FALSE
| Amide Hep120220 T240B1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 72 | 7010_1_021021065.D | Sample |  | 2021-02-11 20:49 |  | 12.1148 | 4361.9577 | FALSE | 11.7008 |  | 11.7008 |  | 59.6192 | FALSE | 11.7881 | FALSE | 10.9972 | 1368939.113 | 16.2463 | FALSE
| Amide Hep120220 T240B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 73 | 7010_1_021021066.D | Sample |  | 2021-02-11 21:22 |  | 12.1148 | 2172.0619 | FALSE | 5.4752 |  | 5.4752 |  |  | FALSE | 13.553 | FALSE | 10.9972 | 1455996.9248 | 16.771 | FALSE
| Amide Hep120220 T240B3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 74 | 7010_1_021021067.D | Sample |  | 2021-02-11 21:55 |  | 12.1148 | 32113.4605 | FALSE | 88.9801 |  | 88.9801 |  | 54.1385 | FALSE | 10.5568 | FALSE | 10.993 | 1334235.7534 | 16.2788 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021068.D | Sample |  | 2021-02-11 22:28 |  | 12.1148 | 19244.6699 | FALSE | 57.1767 |  | 57.1767 |  | 55.842 | FALSE | 5.7351 | FALSE | 10.9972 | 1240868.7448 | 15.9939 | FALSE
| Amide Hep120220 T240B5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 76 | 7010_1_021021069.D | Sample |  | 2021-02-11 23:01 |  | 12.1149 | 8685.8246 | TRUE | 26.5989 |  | 26.5989 |  | 79.2277 | FALSE | 26.8582 | FALSE | 10.993 | 1200683.2968 | 15.7968 | FALSE
| Amide Hep120220 T240B6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 77 | 7010_1_021021070.D | Sample |  | 2021-02-11 23:35 |  | 12.1097 | 6134.533 | FALSE | 21.3198 |  | 21.3198 |  | 52.3185 | FALSE | 8.0814 | FALSE | 10.993 | 1057498.623 | 16.4488 | FALSE
| Amide Hep120220 T0B2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 52 | 7010_1_021021071.D | Sample |  | 2021-02-12 00:08 |  | 12.1148 | 20004.1011 | FALSE | 61.835 |  | 61.835 |  | 53.271 | FALSE | 6.1046 | FALSE | 10.9972 | 1193151.4602 | 16.1777 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021021074.D | Sample |  | 2021-02-12 01:48 |  | 12.1148 | 111923.1403 | FALSE | 269.1499 |  | 269.1499 |  | 50.5928 | FALSE | 7.2764 | FALSE | 10.9972 | 1561871.923 | 16.8299 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021021075.D | Sample |  | 2021-02-12 02:21 |  | 12.1148 | 20892.1979 | FALSE | 62.3533 |  | 62.3533 |  | 52.2129 | FALSE | 6.3734 | FALSE | 10.993 | 1235820.4763 | 17.052 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021021076.D | Sample |  | 2021-02-12 02:54 |  | 12.1097 | 42835.9722 | FALSE | 127.5451 |  | 127.5451 |  | 53.6491 | FALSE | 6.9575 | FALSE | 10.9972 | 1245796.3823 | 17.1506 | FALSE
| Amide Hep120220 T240B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 78 | 7010_1_021021077.D | Sample |  | 2021-02-12 03:27 |  | 12.1097 | 4819.8122 | FALSE | 16.0914 |  | 16.0914 |  | 47.4405 | FALSE | 9.9365 | FALSE | 10.993 | 1100328.3884 | 16.9299 | FALSE
| Amide Hep120220 T240B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 79 | 7010_1_021021078.D | Sample |  | 2021-02-12 04:01 |  | 12.1097 | 4222.2849 | FALSE | 12.92 |  | 12.92 |  | 60.4438 | FALSE | 17.827 | FALSE | 10.993 | 1200195.6816 | 17.4207 | FALSE
| Amide Hep120220 T240B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 80 | 7010_1_021021079.D | Sample |  | 2021-02-12 04:34 |  | 12.1097 | 2906.9116 | FALSE | 9.2504 |  | 9.2504 |  | 61.1759 | FALSE | 19.1033 | FALSE | 10.993 | 1153710.9073 | 17.2878 | FALSE
| Amide Hep120220 T0C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 81 | 7010_1_021021080.D | Sample |  | 2021-02-12 05:07 |  | 12.1097 | 71083.9036 | FALSE | 266.4145 |  | 266.4145 | 47.4477 | 54.0438 | FALSE | 5.5456 | FALSE | 10.993 | 1001908.3613 | 16.7672 | FALSE
| Amide Hep120220 T0C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 82 | 7010_1_021021081.D | Sample |  | 2021-02-12 05:40 |  | 12.1097 | 151238.0407 | FALSE | 601.8854 |  | 601.8854 | 107.1942 | 55.6235 | FALSE | 4.882 | FALSE | 10.993 | 972455.3572 | 16.7866 | FALSE
| Amide Hep120220 T0C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 83 | 7010_1_021021083.D | Sample |  | 2021-02-12 06:47 |  | 12.1097 | 171827.5834 | FALSE | 816.1729 |  | 816.1729 | 145.3582 | 55.2947 | FALSE | 4.6414 | FALSE | 10.993 | 831032.5682 | 16.8694 | FALSE
| Amide Hep120220 T0C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 84 | 7010_1_021021084.D | Sample |  | 2021-02-12 07:20 |  | 12.1097 | 202489.7102 | FALSE | 770.4552 |  | 770.4552 | 112.9468 | 56.4155 | FALSE | 5.0541 | FALSE | 10.993 | 1033039.7387 | 16.4451 | FALSE
| Amide Hep120220 T0C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 85 | 7010_1_021021085.D | Sample |  | 2021-02-12 07:53 |  | 12.1097 | 171143.299 | FALSE | 742.0912 |  | 742.0912 | 108.7887 | 56.2888 | FALSE | 4.7303 | FALSE | 10.993 | 904113.176 | 16.4554 | FALSE
| Amide Hep120220 T0C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 86 | 7010_1_021021086.D | Sample |  | 2021-02-12 08:27 |  | 12.1097 | 129335.8793 | FALSE | 533.8737 |  | 533.8737 | 78.2645 | 55.8344 | FALSE | 5.884 | FALSE | 10.993 | 931779.9049 | 16.8305 | FALSE
| Amide Hep120220 T0C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 87 | 7010_1_021021089.D | Sample |  | 2021-02-12 10:06 |  | 12.1097 | 88986.4844 | FALSE | 370.3071 |  | 370.3071 | 77.8272 | 55.9451 | FALSE | 5.0319 | FALSE | 10.993 | 910738.7204 | 16.7235 | FALSE
| Amide Hep120220 T0C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 88 | 7010_1_021021090.D | Sample |  | 2021-02-12 10:39 |  | 12.1097 | 262005.5584 | FALSE | 672.6962 |  | 672.6962 | 141.3801 | 53.7135 | FALSE | 6.0404 | FALSE | 10.993 | 1517162.4518 | 16.8494 | FALSE
| Amide Hep120220 T0C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 89 | 7010_1_021021091.D | Sample |  | 2021-02-12 11:21 |  | 12.1148 | 136813.7205 | FALSE | 384.4173 |  | 384.4173 | 80.7927 | 53.169 | FALSE | 6.8471 | FALSE | 10.993 | 1350538.834 | 16.6547 | FALSE
| Amide Hep120220 T240B4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 75 | 7010_1_021021110.D | Sample |  | 2021-02-12 21:53 |  | 12.1046 | 20474.5629 | FALSE | 58.2249 |  | 58.2249 |  | 53.5253 | FALSE | 9.4849 | FALSE | 10.9887 | 1296524.3149 | 16.4413 | FALSE
| Amide Hep120220 T240C1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 102 | 7010_1_021021111.D | Sample |  | 2021-02-12 22:26 |  | 12.1046 | 35512.4533 | FALSE | 124.0698 |  | 124.0698 | 131.9618 | 49.1441 | FALSE | 14.5943 | FALSE | 10.993 | 1061413.605 | 16.4233 | FALSE
| Amide Hep120220 T240C2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 103 | 7010_1_021021112.D | Sample |  | 2021-02-12 22:59 |  | 12.1097 | 54138.6211 | FALSE | 68.9066 |  | 68.9066 | 73.2897 | 47.8768 | FALSE | 13.2013 | FALSE | 10.993 | 2899512.1558 | 17.9619 | FALSE
| Amide Hep120220 T240C3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 104 | 7010_1_021021113.D | Sample |  | 2021-02-12 23:33 |  | 12.1097 | 72245.6513 | FALSE | 89.0821 |  | 89.0821 | 94.7486 | 48.3554 | FALSE | 15.8465 | FALSE | 10.993 | 2998218.9334 | 17.8928 | FALSE
| Amide Hep120220 T240C4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 105 | 7010_1_021021114.D | Sample |  | 2021-02-13 00:06 |  | 12.1046 | 116196.575 | FALSE | 181.8669 |  | 181.8669 | 42.8825 | 42.8284 | FALSE | 12.8613 | FALSE | 10.9887 | 2381288.7606 | 17.9741 | FALSE
| Amide Hep120220 T240C5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 106 | 7010_1_021021115.D | Sample |  | 2021-02-13 00:39 |  | 12.1097 | 133853.1324 | FALSE | 379.2173 |  | 379.2173 | 89.416 | 53.1709 | FALSE | 8.8182 | FALSE | 10.9887 | 1338808.5188 | 16.8042 | FALSE
| Amide Hep120220 T240C6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 107 | 7010_1_021021116.D | Sample |  | 2021-02-13 01:12 |  | 12.1097 | 261398.3489 | FALSE | 711.23 |  | 711.23 | 167.7015 | 52.6528 | FALSE | 8.1553 | FALSE | 10.9887 | 1436728.2402 | 16.4115 | FALSE
| Amide Hep120220 T240C7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 108 | 7010_1_021021117.D | Sample |  | 2021-02-13 01:46 |  | 12.1046 | 247868.5832 | FALSE | 665.7928 |  | 665.7928 | 102.7957 | 50.2309 | FALSE | 10.4379 | FALSE | 10.9887 | 1449263.6916 | 16.4455 | FALSE
| Amide Hep120220 T240C8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 109 | 7010_1_021021118.D | Sample |  | 2021-02-13 02:19 |  | 12.1097 | 178177.9164 | FALSE | 567.4671 |  | 567.4671 | 87.6146 | 53.8051 | FALSE | 7.4924 | FALSE | 10.9887 | 1211357.7857 | 16.428 | FALSE
| Amide Hep120220 T240C9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 110 | 7010_1_021021119.D | Sample |  | 2021-02-13 02:52 |  | 12.1046 | 177061.6147 | FALSE | 709.7973 |  | 709.7973 | 109.5898 | 53.195 | FALSE | 5.9039 | FALSE | 10.993 | 975022.2483 | 15.7866 | FALSE
| Amide Hep120220 T0B7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 57 | 7010_1_021521004.D | Sample |  | 2021-02-15 11:17 |  | 12.1098 | 111897.932 | FALSE | 294.9761 |  | 294.9761 |  | 54.6607 | FALSE | 5.8676 | FALSE | 10.9888 | 1428073.1821 | 16.0304 | FALSE
| Amide Hep120220 T0B8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 58 | 7010_1_021521005.D | Sample |  | 2021-02-15 11:50 |  | 12.1047 | 24859.1538 | FALSE | 71.1038 |  | 71.1038 |  | 47.8617 | FALSE | 10.1363 | FALSE | 10.9888 | 1290491.2495 | 15.8517 | FALSE
| Amide Hep120220 T0B9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 59 | 7010_1_021521006.D | Sample |  | 2021-02-15 12:24 |  | 12.1098 | 49862.9396 | FALSE | 154.905 |  | 154.905 |  | 54.8262 | FALSE | 6.6011 | FALSE | 10.9888 | 1196894.7839 | 16.3618 | FALSE
| Amide Hep120220 TMidB1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 60 | 7010_1_021521007.D | Sample |  | 2021-02-15 12:57 |  | 12.1047 | 33730.7861 | FALSE | 92.2789 |  | 92.2789 |  | 47.7647 | FALSE | 8.1302 | FALSE | 10.9888 | 1351722.6553 | 16.3029 | FALSE
| Amide Hep120220 TMidB2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 61 | 7010_1_021521008.D | Sample |  | 2021-02-15 13:30 |  | 12.1047 | 53042.303 | FALSE | 132.7787 |  | 132.7787 |  | 52.9782 | FALSE | 9.2573 | FALSE | 10.9888 | 1482501.358 | 15.7608 | FALSE
| Amide Hep120220 TMidB3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 62 | 7010_1_021521009.D | Sample |  | 2021-02-15 14:03 |  | 12.1047 | 17350.4875 | FALSE | 40.3338 |  | 40.3338 |  | 63.8247 | FALSE | 6.0908 | FALSE | 10.9889 | 1583586.7699 | 16.5115 | FALSE
| Amide Hep120220 TMidB4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 63 | 7010_1_021521010.D | Sample |  | 2021-02-15 14:37 |  | 12.1047 | 12949.6496 | FALSE | 34.6316 |  | 34.6316 |  | 52.2058 | FALSE | 6.8423 | FALSE | 10.9889 | 1375845.0535 | 15.8883 | FALSE
| Amide Hep120220 TMidB5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 64 | 7010_1_021521011.D | Sample |  | 2021-02-15 15:10 |  | 12.1047 | 26983.0177 | FALSE | 66.7336 |  | 66.7336 |  | 51.3811 | FALSE | 6.3019 | FALSE | 10.9888 | 1491909.1765 | 16.5452 | FALSE
| Amide Hep120220 TMidB6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 65 | 7010_1_021521012.D | Sample |  | 2021-02-15 15:43 |  | 12.1047 | 8181.9628 | FALSE | 21.7066 |  | 21.7066 |  | 49.8694 | FALSE | 4.0173 | FALSE | 10.9888 | 1385355.4716 | 16.0732 | FALSE
| Amide Hep120220 TMidB7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 66 | 7010_1_021521013.D | Sample |  | 2021-02-15 16:17 |  | 12.1047 | 4156.9177 | FALSE | 10.1623 |  | 10.1623 |  | 60.7395 | FALSE | 4.9699 | FALSE | 10.9888 | 1501903.8251 | 15.9211 | FALSE
| Amide Hep120220 TMidB8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 67 | 7010_1_021521016.D | Sample |  | 2021-02-15 17:56 |  | 12.1047 | 23205.1124 | FALSE | 55.9058 |  | 55.9058 |  | 60.2353 | FALSE | 5.9788 | FALSE | 10.9846 | 1530077.5364 | 16.7428 | FALSE
| Amide Hep120220 TMidB9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 68 | 7010_1_021521017.D | Sample |  | 2021-02-15 18:30 |  | 12.1047 | 33473.7145 | FALSE | 79.0292 |  | 79.0292 |  | 50.6837 | FALSE | 7.8207 | FALSE | 10.9846 | 1564508.1695 | 16.4263 | FALSE
| Amide Hep120220 TMidB10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 69 | 7010_1_021521018.D | Sample |  | 2021-02-15 19:03 |  | 12.1047 | 18178.9767 | FALSE | 42.058 |  | 42.058 |  | 55.0422 | FALSE | 9.9078 | FALSE | 10.9889 | 1591418.6379 | 16.5783 | FALSE
| Amide Hep120220 TMidB11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 70 | 7010_1_021521019.D | Sample |  | 2021-02-15 19:36 |  | 12.1047 | 13882.4397 | FALSE | 32.8208 |  | 32.8208 |  | 57.6702 | FALSE | 16.1725 | FALSE | 10.9888 | 1556081.2329 | 16.1663 | FALSE
| Amide Hep120220 TMidB12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 71 | 7010_1_021521020.D | Sample |  | 2021-02-15 20:09 |  | 12.1047 | 22049.7862 | FALSE | 74.7526 |  | 74.7526 |  | 54.944 | FALSE | 7.5394 | FALSE | 10.9846 | 1089125.1205 | 15.782 | FALSE
| Amide Hep120220 TMidC1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 90 | 7010_1_021521022.D | Sample |  | 2021-02-15 21:16 |  | 12.1047 | 177997.6651 | FALSE | 510.7829 |  | 510.7829 | 105.3049 | 54.914 | FALSE | 7.8741 | FALSE | 10.9846 | 1337523.5952 | 15.7728 | FALSE
| Amide Hep120220 TMidC2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 91 | 7010_1_021521023.D | Sample |  | 2021-02-15 21:49 |  | 12.1047 | 73494.0659 | FALSE | 211.0287 |  | 211.0287 | 43.5064 | 55.7816 | FALSE | 7.9159 | FALSE | 10.9846 | 1301363.9414 | 15.9555 | FALSE
| Amide Hep120220 TMidC3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 92 | 7010_1_021521024.D | Sample |  | 2021-02-15 22:22 |  | 12.1047 | 234244.8623 | FALSE | 733.3433 |  | 733.3433 | 151.1887 | 55.787 | FALSE | 5.2102 | FALSE | 10.9846 | 1251213.9398 | 15.6121 | FALSE
| Amide Hep120220 TMidC4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 93 | 7010_1_021521025.D | Sample |  | 2021-02-15 22:56 |  | 12.1047 | 173449.5263 | FALSE | 668.5907 |  | 668.5907 | 127.4182 | 52.8 | FALSE | 7.2547 | FALSE | 10.9846 | 1010158.4132 | 15.6893 | FALSE
| Amide Hep120220 TMidC5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 94 | 7010_1_021521026.D | Sample |  | 2021-02-15 23:29 |  | 12.1047 | 114145.4618 | FALSE | 448.468 |  | 448.468 | 85.4678 | 57.2819 | FALSE | 6.0112 | FALSE | 10.9846 | 971417.7653 | 15.9506 | FALSE
| Amide Hep120220 TMidC6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 95 | 7010_1_021521027.D | Sample |  | 2021-02-16 00:02 |  | 12.1047 | 162561.8763 | FALSE | 457.1061 |  | 457.1061 | 87.114 | 54.4828 | FALSE | 5.0906 | FALSE | 10.9846 | 1358371.3736 | 16.5259 | FALSE
| Amide Hep120220 TMidC7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 96 | 7010_1_021521030.D | Sample |  | 2021-02-16 01:42 |  | 12.1047 | 265351.9561 | FALSE | 757.0915 |  | 757.0915 | 142.5413 | 52.1611 | FALSE | 5.1815 | FALSE | 10.9337 | 1375933.0656 | 16.6485 | TRUE
| Amide Hep120220 TMidC8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 97 | 7010_1_021521031.D | Sample |  | 2021-02-16 02:15 |  | 12.1047 | 148696.2105 | FALSE | 399.0999 |  | 399.0999 | 75.1405 | 55.8194 | FALSE | 4.3205 | FALSE | 10.9888 | 1415696.2804 | 16.9998 | FALSE
| Amide Hep120220 TMidC9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 98 | 7010_1_021521032.D | Sample |  | 2021-02-16 02:48 |  | 12.1047 | 134972.029 | FALSE | 437.2235 |  | 437.2235 | 82.3182 | 52.0211 | FALSE | 5.7138 | FALSE | 10.9846 | 1177007.9129 | 16.2408 | FALSE
| Amide Hep120220 TMidC10_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 99 | 7010_1_021521033.D | Sample |  | 2021-02-16 03:22 |  | 12.1047 | 96979.5995 | FALSE | 230.7708 |  | 230.7708 | 118.8476 | 54.7978 | FALSE | 5.5431 | FALSE | 10.9846 | 1573056.5806 | 16.6627 | FALSE
| Amide Hep120220 TMidC11_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 100 | 7010_1_021521034.D | Sample |  | 2021-02-16 03:55 |  | 12.1047 | 124144.6916 | FALSE | 314.8108 |  | 314.8108 | 162.1285 | 54.0495 | FALSE | 5.1253 | FALSE | 10.9846 | 1487167.1693 | 16.3383 | FALSE
| Amide Hep120220 TMidC12_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 101 | 7010_1_021521035.D | Sample |  | 2021-02-16 04:28 |  | 12.1047 | 72560.4933 | FALSE | 222.2775 |  | 222.2775 | 114.4736 | 55.2403 | FALSE | 4.125 | FALSE | 10.9846 | 1221022.527 | 16.1578 | FALSE
| Amide Hep120220 T240A1_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 42 | 7010_1_021521036.D | Sample |  | 2021-02-16 05:01 |  | 12.1047 | 3504.0906 | FALSE | 8.8356 |  | 8.8356 |  | 61.3272 | FALSE | 2.4032 | FALSE | 10.9846 | 1455972.8968 | 16.8817 | FALSE
| Amide Hep120220 T240A2_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 43 | 7010_1_021521037.D | Sample |  | 2021-02-16 05:35 |  | 12.0996 | 1278.6248 | FALSE | 4.1143 |  | 4.1143 |  | 63.9379 | FALSE | 7.6304 | FALSE | 10.9846 | 1140475.4058 | 16.5723 | FALSE
| Amide Hep120220 T240A3_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 44 | 7010_1_021521038.D | Sample |  | 2021-02-16 06:08 |  | 12.0996 | 10514.8557 | FALSE | 35.041 |  | 35.041 |  | 50.6125 | FALSE | 4.5498 | FALSE | 10.9846 | 1104145.2921 | 16.3082 | FALSE
| Amide Hep120220 T240A4_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 45 | 7010_1_021521040.D | Sample |  | 2021-02-16 07:15 |  | 12.0996 | 11559.8461 | FALSE | 39.9244 |  | 39.9244 |  | 61.6172 | FALSE | 5.4107 | FALSE | 10.9846 | 1065852.6822 | 16.3473 | FALSE
| Amide Hep120220 T240A5_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 46 | 7010_1_021521041.D | Sample |  | 2021-02-16 07:48 |  | 12.0996 | 6021.1338 | FALSE | 23.1586 |  | 23.1586 |  | 56.9541 | FALSE | 5.283 | FALSE | 10.9846 | 955690.2177 | 16.0574 | FALSE
| Amide Hep120220 T240A6_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 47 | 7010_1_021521042.D | Sample |  | 2021-02-16 08:21 |  | 12.0997 | 2102.2594 | FALSE | 8.2845 |  | 8.2845 |  | 103.124 | FALSE | 5.7796 | FALSE | 10.9846 | 931565.1466 | 16.0569 | FALSE
| Amide Hep120220 T240A7_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 48 | 7010_1_021521043.D | Sample |  | 2021-02-16 08:54 |  | 12.1047 | 3231.7855 | FALSE | 7.7697 |  | 7.7697 |  | 58.6644 | FALSE | 4.1103 | FALSE | 10.9804 | 1526895.0041 | 15.9114 | FALSE
| Amide Hep120220 T240A8_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 49 | 7010_1_021521045.D | Sample |  | 2021-02-16 10:01 |  | 12.1047 | 2102.175 | TRUE | 7.3797 |  | 7.3797 |  | 41.6135 | TRUE | 11.2381 | TRUE | 10.9804 | 1045655.2677 | 16.0693 | FALSE
| Amide Hep120220 T240A9_020921 | E:\Shark Tank\Hep\021021 | HC Amides 12/2 | 50 | 7010_1_021521046.D | Sample |  | 2021-02-16 10:34 |  | 12.0945 | 1070.3122 | TRUE | 3.6022 |  | 3.6022 |  | 34.2888 | TRUE |  | FALSE | 10.9804 | 1090324.8465 | 16.5326 | FALSE

## 923
Sample | [8 empty] | 923 Method | 923 Results | [6 empty] | Qualifier (224.9 -> 44.1) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | RT | Resp. | Ratio | MI
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/30 | 71 | 7010_1_033021001.D | MatrixBlank |  | 2021-03-30 10:21 |  | 24.3472 | 533.1367 | FALSE | 0.74909 |  | 0.74909 |  | 1082.9804 | FALSE | 10.8984 | 4601677.409 | 57.4176 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/31 | 71 | 7010_1_033021002.D | MatrixBlank |  | 2021-03-30 10:54 |  | 24.288 | 534.9544 | FALSE | 0.76764 |  | 0.76764 |  | 1152.9383 | FALSE | 10.8891 | 4505772.0892 | 60.0572 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/32 | 71 | 7010_1_033021003.D | MatrixBlank |  | 2021-03-30 11:27 |  | 24.2541 | 211.4347 | FALSE | 0.28749 |  | 0.28749 |  | 1572.4892 | FALSE | 10.8891 | 4755573.0829 | 58.2263 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/33 | 71 | 7010_1_033021004.D | MatrixBlank |  | 2021-03-30 12:00 |  | 24.1695 | 214.1245 | FALSE | 0.30275 |  | 0.30275 |  |  | FALSE | 10.8891 | 4573351.7197 | 58.3867 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/34 | 71 | 7010_1_033021005.D | MatrixBlank |  | 2021-03-30 12:33 |  | 23.9411 | 267.3624 | FALSE | 0.36113 |  | 0.36113 |  | 12093.7241 | FALSE | 10.8891 | 4787098.2396 | 58.9397 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/35 | 72 | 7010_1_033021006.D | QC | 15 | 2021-03-30 13:07 | 2000 | 24.3049 | 2081137.0026 | FALSE | 1903.7975 |  | 1903.7975 | 95.1899 | 605.7031 | FALSE | 10.8891 | 5308432.1467 | 56.7064 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/36 | 72 | 7010_1_033021007.D | QC | 15 | 2021-03-30 13:40 | 2000 | 24.2965 | 2000178.7561 | FALSE | 1872.5294 |  | 1872.5294 | 93.6265 | 598.8982 | FALSE | 10.8891 | 5208426.2116 | 57.9386 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/37 | 72 | 7010_1_033021008.D | QC | 15 | 2021-03-30 14:13 | 2000 | 24.288 | 1876875.7018 | FALSE | 1845.0927 |  | 1845.0927 | 92.2546 | 601.3171 | FALSE | 10.8891 | 4977962.3769 | 59.1409 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/38 | 72 | 7010_1_033021009.D | QC | 15 | 2021-03-30 14:46 | 2000 | 24.2965 | 2149838.8476 | FALSE | 2072.4182 |  | 2072.4182 | 103.6209 | 603.5718 | FALSE | 10.8891 | 4928781.2375 | 60.1901 | FALSE
| HC_Amide CC 15 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/39 | 72 | 7010_1_033021010.D | Cal | 15 | 2021-03-30 15:19 | 2000 | 24.288 | 2155642.8498 | FALSE | 1978.4571 |  | 1978.4571 | 98.9229 | 604.131 | FALSE | 10.8891 | 5239811.2713 | 57.4166 | FALSE
| HC_Amide CC 14 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/40 | 73 | 7010_1_033021011.D | Cal | 14 | 2021-03-30 15:53 | 1250 | 24.288 | 1478876.8057 | FALSE | 1370.5648 |  | 1370.5648 | 109.6452 | 607.7799 | FALSE | 10.8891 | 5632749.417 | 56.1521 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/41 | 74 | 7010_1_033021012.D | Cal | 13 | 2021-03-30 16:26 | 781 | 24.288 | 798408.9969 | FALSE | 808.5304 |  | 808.5304 | 103.525 | 613.6945 | FALSE | 10.8891 | 5597229.1939 | 56.1416 | FALSE
| HC_Amide CC 12 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/42 | 75 | 7010_1_033021013.D | Cal | 12 | 2021-03-30 16:59 | 488 | 24.2965 | 487042.3229 | FALSE | 524.2495 |  | 524.2495 | 107.4282 | 608.3834 | FALSE | 10.8891 | 5504844.517 | 57.0132 | FALSE
| HC_Amide CC 11 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/43 | 76 | 7010_1_033021014.D | Cal | 11 | 2021-03-30 17:32 | 305 | 24.2795 | 311628.0282 | FALSE | 341.8474 |  | 341.8474 | 112.0811 | 606.2102 | FALSE | 10.8891 | 5563556.708 | 57.1378 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/44 | 77 | 7010_1_033021015.D | Cal | 10 | 2021-03-30 18:06 | 191 | 24.2795 | 188189.4612 | FALSE | 212.2565 |  | 212.2565 | 111.1291 | 622.8713 | FALSE | 10.8891 | 5528855.0073 | 57.8612 | FALSE
| HC_Amide CC 9 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/45 | 78 | 7010_1_033021016.D | Cal | 9 | 2021-03-30 18:39 | 119 | 24.2795 | 123406.6549 | FALSE | 145.713 |  | 145.713 | 122.4479 | 594.486 | FALSE | 10.8891 | 5341005.6985 | 58.3686 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/46 | 79 | 7010_1_033021017.D | Cal | 8 | 2021-03-30 19:12 | 75 | 24.2795 | 83642.6005 | FALSE | 95.5594 |  | 95.5594 | 127.4125 | 602.2875 | FALSE | 10.8891 | 5567409.6743 | 57.6535 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/47 | 80 | 7010_1_033021018.D | Cal | 7 | 2021-03-30 19:46 | 47 | 24.2711 | 53826.0251 | FALSE | 60.0183 |  | 60.0183 | 127.6984 | 602.7263 | FALSE | 10.8891 | 5739321.8892 | 57.2117 | FALSE
| HC_Amide CC 6 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/48 | 81 | 7010_1_033021019.D | Cal | 6 | 2021-03-30 20:19 | 29 | 24.2711 | 33614.1814 | FALSE | 36.6968 |  | 36.6968 | 126.5406 | 606.9524 | FALSE | 10.8891 | 5885668.5836 | 56.5328 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/49 | 82 | 7010_1_033021020.D | Cal | 5 | 2021-03-30 20:52 | 18 | 24.2711 | 21041.3343 | FALSE | 23.0313 |  | 23.0313 | 127.9515 | 609.3099 | FALSE | 10.8891 | 5884165.0948 | 56.941 | FALSE
| HC_Amide CC 4 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/50 | 83 | 7010_1_033021021.D | Cal | 4 | 2021-03-30 21:25 | 11 | 24.2626 | 12058.3869 | FALSE | 13.2253 |  | 13.2253 | 120.2298 | 570.84 | FALSE | 10.8891 | 5882387.0073 | 56.8245 | FALSE
| HC_Amide CC 3 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/51 | 84 | 7010_1_033021022.D | Cal | 3 | 2021-03-30 21:59 | 7.1 | 24.2457 | 7529.1178 | FALSE | 8.5288 |  | 8.5288 | 120.1243 | 649.0889 | FALSE | 10.8891 | 5700048.6818 | 57.1476 | FALSE
| HC_Amide CC 2 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/52 | 85 | 7010_1_033021023.D | Cal | 2 | 2021-03-30 22:32 | 4.4 | 24.2626 | 4377.39 | FALSE | 4.8004 |  | 4.8004 | 109.0994 | 564.4809 | FALSE | 10.8797 | 5891768.6437 | 56.6901 | FALSE
| HC_Amide CC 1 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/53 | 86 | 7010_1_033021024.D | Cal | 1 | 2021-03-30 23:05 | 2.8 | 24.2457 | 2720.8276 | FALSE | 2.9643 |  | 2.9643 | 105.8674 | 542.4126 | FALSE | 10.8891 | 5932324.5941 | 56.9682 | FALSE
| HC_Amide QCCC 3 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/54 | 87 | 7010_1_033021025.D | QC | 3 | 2021-03-30 23:38 | 7.1 | 24.2288 | 3174.4287 | FALSE | 3.1571 |  | 3.1571 | 44.4665 | 579.8219 | FALSE | 10.8891 | 6498368.3403 | 56.094 | FALSE
| HC_Amide CC 10 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/65 | 77 | 7010_1_033021036.D | Cal | 10 | 2021-03-31 05:44 | 191 | 24.2372 | 136331.1543 | FALSE | 161.7985 |  | 161.7985 | 84.7113 | 602.7306 | FALSE | 10.8797 | 5299295.7285 | 59.7302 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/66 | 89 | 7010_1_033021037.D | QC | 14 | 2021-03-31 06:17 | 1250 | 24.2288 | 491214.3724 | FALSE | 440.4337 |  | 440.4337 | 35.2347 | 609.4408 | FALSE | 10.8797 | 6698171.429 | 57.406 | FALSE
| HC_Amide QCCC 11 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/73 | 88 | 7010_1_033021044.D | QC | 11 | 2021-03-31 10:10 | 305 | 24.2203 | 94622.9561 | FALSE | 92.9494 |  | 92.9494 | 30.4752 | 612.6027 | FALSE | 10.8797 | 6478031.2414 | 54.9648 | FALSE
| HC_Amide CC 7 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/47 | 80 | 7010_1_033021045.D | Cal | 7 | 2021-03-31 10:43 | 47 | 24.2118 | 44224.2424 | FALSE | 47.3715 |  | 47.3715 | 100.7904 | 609.505 | FALSE | 10.8797 | 5987464.417 | 55.9204 | FALSE
| HC_Amide CC 13 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/41 | 74 | 7010_1_033021055.D | Cal | 13 | 2021-03-31 16:16 | 781 | 24.1949 | 544141.0481 | FALSE | 563.0981 |  | 563.0981 | 72.0996 | 600.7645 | FALSE | 10.8797 | 5690614.455 | 57.3219 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/30 | 71 | 7010_1_033021056.D | MatrixBlank |  | 2021-03-31 16:49 |  | 24.2034 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.8797 | 5617561.8075 | 56.6595 | FALSE
| HC_Amide CC 8 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/46 | 79 | 7010_1_033021066.D | Cal | 8 | 2021-04-01 14:32 | 75 | 24.178 | 48546.3861 | FALSE | 54.5589 |  | 54.5589 | 72.7452 | 624.6393 | FALSE | 10.8797 | 5699698.2287 | 58.2552 | FALSE
| HC_Amide QCCC 14 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/124 | 89 | 7010_1_033021071.D | QC | 14 | 2021-04-01 17:18 | 1250 | 24.1695 | 275370.048 | FALSE | 249.3755 |  | 249.3755 | 19.95 | 599.8198 | FALSE | 10.8704 | 6843277.9525 | 58.5758 | FALSE
| Method Spike | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/30 | 71 | 7010_1_033021072.D | MatrixBlank |  | 2021-04-01 17:51 |  | 24.178 | 0 | TRUE | 0 |  | 0 |  |  | FALSE | 10.8797 | 5478148.5023 | 59.5837 | FALSE
| HC_Amide CC 5 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/49 | 82 | 7010_1_033021073.D | Cal | 5 | 2021-04-01 18:25 | 18 | 24.1695 | 11419.5542 | FALSE | 14.5171 |  | 14.5171 | 80.6504 | 555.0944 | FALSE | 10.8704 | 5073897.4886 | 61.1146 | FALSE
| Amide Hep120220 T0D1_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/55 | 90 | 7010_1_033021026.D | Sample |  | 2021-03-31 00:12 |  | 24.2541 | 282748.0926 | FALSE | 263.3388 |  | 263.3388 | 114.3531 | 603.8709 | FALSE | 10.8611 | 6638578.0779 | 56.7008 | FALSE
| Amide Hep120220 T0D2_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/56 | 91 | 7010_1_033021027.D | Sample |  | 2021-03-31 00:45 |  | 24.2626 | 230694.6858 | FALSE | 217.7007 |  | 217.7007 | 94.5351 | 614.9271 | FALSE | 10.8797 | 6602091.8045 | 56.0978 | FALSE
| Amide Hep120220 T0D3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/57 | 92 | 7010_1_033021028.D | Sample |  | 2021-03-31 01:18 |  | 24.2457 | 225598.4087 | FALSE | 209.8174 |  | 209.8174 | 91.1118 | 617.1379 | FALSE | 10.8704 | 6707695.175 | 56.3865 | FALSE
| Amide Hep120220 T0D3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/58 | 92 | 7010_1_033021029.D | Sample |  | 2021-03-31 01:51 |  | 24.2457 | 225505.2952 | FALSE | 217.8155 |  | 217.8155 |  | 608.7981 | FALSE | 10.8797 | 6450056.7749 | 56.0379 | FALSE
| Amide Hep120220 T0D3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/59 | 92 | 7010_1_033021030.D | Sample |  | 2021-03-31 02:25 |  | 24.2541 | 219330.0849 | FALSE | 213.9934 |  | 213.9934 |  | 611.3128 | FALSE | 10.8797 | 6389575.6126 | 55.7986 | FALSE
| Amide Hep120220 T0D4_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/60 | 93 | 7010_1_033021031.D | Sample |  | 2021-03-31 02:58 |  | 24.2457 | 220094.1255 | FALSE | 209.1715 |  | 209.1715 | 100.67 | 620.4516 | FALSE | 10.8611 | 6564958.2786 | 56.6464 | FALSE
| Amide Hep120220 T0D5_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/61 | 94 | 7010_1_033021032.D | Sample |  | 2021-03-31 03:31 |  | 24.2541 | 211688.5889 | FALSE | 207.5994 |  | 207.5994 | 99.9134 | 626.2153 | FALSE | 10.8611 | 6363735.721 | 56.9601 | FALSE
| Amide Hep120220 T0D6_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/62 | 95 | 7010_1_033021033.D | Sample |  | 2021-03-31 04:04 |  | 24.2288 | 214164.0617 | FALSE | 206.5671 |  | 206.5671 | 99.4166 | 614.0062 | FALSE | 10.8611 | 6471448.6166 | 56.4873 | FALSE
| Amide Hep120220 T0D7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/63 | 96 | 7010_1_033021034.D | Sample |  | 2021-03-31 04:38 |  | 24.2372 | 213041.6738 | FALSE | 214.7093 |  | 214.7093 | 94.7512 | 620.0025 | FALSE | 10.8611 | 6184942.0302 | 57.5748 | FALSE
| Amide Hep120220 T0D7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/64 | 96 | 7010_1_033021035.D | Sample |  | 2021-03-31 05:11 |  | 24.2372 | 223305.2022 | FALSE | 224.4136 |  | 224.4136 |  | 607.6848 | FALSE | 10.8611 | 6192478.8312 | 57.1991 | FALSE
| Amide Hep120220 T0D8_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/67 | 97 | 7010_1_033021038.D | Sample |  | 2021-03-31 06:51 |  | 24.2372 | 245054.6666 | FALSE | 240.6869 |  | 240.6869 | 106.2151 | 617.6113 | FALSE | 10.8611 | 6318912.561 | 57.2527 | FALSE
| Amide Hep120220 T0D9_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/68 | 98 | 7010_1_033021039.D | Sample |  | 2021-03-31 07:24 |  | 24.2372 | 229955.0832 | FALSE | 228.6185 |  | 228.6185 | 100.8893 | 604.4612 | FALSE | 10.8611 | 6255191.3343 | 57.4124 | FALSE
| Amide Hep120220 TMidD1_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/69 | 99 | 7010_1_033021040.D | Sample | 15 | 2021-03-31 07:57 |  | 24.2288 | 301222.0973 | FALSE | 291.9118 |  | 291.9118 | 94.0242 | 616.5594 | FALSE | 10.8718 | 6349850.956 | 56.9086 | FALSE
| Amide Hep120220 TMidD2_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/70 | 100 | 7010_1_033021041.D | Sample | 12 | 2021-03-31 08:30 |  | 24.2288 | 236208.1989 | FALSE | 225.036 |  | 225.036 | 72.4837 | 615.4489 | FALSE | 10.8797 | 6531493.6499 | 56.5428 | FALSE
| Amide Hep120220 TMidD3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/71 | 101 | 7010_1_033021042.D | Sample | 11 | 2021-03-31 09:04 |  | 24.2288 | 451360.1981 | FALSE | 414.4454 |  | 414.4454 | 133.4921 | 614.3065 | FALSE | 10.8797 | 6568276.8412 | 56.3102 | FALSE
| Amide Hep120220 T0D3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/72 | 92 | 7010_1_033021043.D | Sample |  | 2021-03-31 09:37 |  | 24.2288 | 198741.2341 | FALSE | 195.0589 |  | 195.0589 |  | 611.8294 | FALSE | 10.8797 | 6372056.5435 | 56.5728 | FALSE
| Amide Hep120220 T0D7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/64 | 96 | 7010_1_033021046.D | Sample |  | 2021-03-31 11:17 |  | 24.2118 | 205487.3288 | FALSE | 207.3043 |  | 207.3043 |  | 607.0509 | FALSE | 10.8611 | 6186415.9574 | 56.957 | FALSE
| Amide Hep120220 TMidD4_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/76 | 102 | 7010_1_033021047.D | Sample | 10 | 2021-03-31 11:50 |  | 24.2118 | 422705.6146 | FALSE | 391.931 |  | 391.931 | 112.7812 | 613.6112 | FALSE | 10.8718 | 6528529.8495 | 57.1893 | FALSE
| Amide Hep120220 TMidD5_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/77 | 103 | 7010_1_033021048.D | Sample | 9 | 2021-03-31 12:23 |  | 24.2118 | 328505.6159 | FALSE | 314.083 |  | 314.083 | 90.3799 | 615.9999 | FALSE | 10.8704 | 6412590.2955 | 56.7233 | FALSE
| Amide Hep120220 TMidD6_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/78 | 104 | 7010_1_033021049.D | Sample | 8 | 2021-03-31 12:56 |  | 24.2034 | 350599.5908 | FALSE | 336.5291 |  | 336.5291 | 112.7812 | 612.9375 | FALSE | 10.8797 | 6363806.4514 | 56.5452 | FALSE
| Amide Hep120220 TMidD6_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/79 | 104 | 7010_1_033021050.D | Sample | 8 | 2021-03-31 13:29 |  | 24.2203 | 382324.7195 | FALSE | 368.0777 |  | 368.0777 |  | 609.3027 | FALSE | 10.8797 | 6312078.3683 | 57.2655 | FALSE
| Amide Hep120220 TMidD7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/80 | 105 | 7010_1_033021051.D | Sample | 7 | 2021-03-31 14:03 |  | 24.2034 | 458617.4066 | FALSE | 437.1749 |  | 437.1749 | 109.5382 | 619.9851 | FALSE | 10.8688 | 6303620.1644 | 57.875 | FALSE
| Amide Hep120220 TMidD8_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/81 | 106 | 7010_1_033021052.D | Sample | 6 | 2021-03-31 14:36 |  | 24.2034 | 467428.9968 | FALSE | 457.9483 |  | 457.9483 | 114.7432 | 609.5147 | FALSE | 10.8683 | 6112742.279 | 57.9656 | FALSE
| Amide Hep120220 TMidD9_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/82 | 107 | 7010_1_033021053.D | Sample | 5 | 2021-03-31 15:09 |  | 24.2034 | 311139.4729 | FALSE | 302.1986 |  | 302.1986 | 75.7186 | 621.1381 | FALSE | 10.8797 | 6324861.1936 | 57.5698 | FALSE
| Amide Hep120220 TMidD10_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/83 | 108 | 7010_1_033021054.D | Sample | 4 | 2021-03-31 15:42 |  | 24.1949 | 301374.8288 | FALSE | 293.3945 |  | 293.3945 | 96.6982 | 605.2457 | FALSE | 10.8704 | 6319411.7965 | 57.2266 | FALSE
| Amide Hep120220 TMidD11_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/87 | 109 | 7010_1_033021057.D | Sample | 4 | 2021-04-01 09:33 |  | 24.2287 | 324182.4172 | FALSE | 338.7989 |  | 338.7989 | 111.6627 | 602.0409 | FALSE | 10.8712 | 5842697.9082 | 58.3714 | FALSE
| Amide Hep120220 TMidD12_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/88 | 110 | 7010_1_033021058.D | Sample | 4 | 2021-04-01 10:07 |  | 24.2034 | 300152.7601 | FALSE | 278.0447 |  | 278.0447 | 91.6391 | 616.4159 | FALSE | 10.8517 | 6658180.3978 | 57.3577 | FALSE
| Amide Hep120220 TMidD6_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/79 | 104 | 7010_1_033021059.D | Sample | 8 | 2021-04-01 10:40 |  | 24.1949 | 326654.2544 | FALSE | 313.1483 |  | 313.1483 |  | 615.5749 | FALSE | 10.8704 | 6396470.2343 | 57.5727 | FALSE
| Amide Hep120220 T240D1_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/90 | 111 | 7010_1_033021060.D | Sample |  | 2021-04-01 11:13 |  | 24.1865 | 236861.2284 | FALSE | 231.3611 |  | 231.3611 | 143.9529 | 615.3263 | FALSE | 10.8657 | 6363749.7023 | 57.3594 | FALSE
| Amide Hep120220 T240D2_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/91 | 112 | 7010_1_033021061.D | Sample |  | 2021-04-01 11:46 |  | 24.1949 | 105421.9977 | FALSE | 108.0965 |  | 108.0965 | 106.3556 | 598.32 | FALSE | 10.8704 | 6189944.3317 | 57.6842 | FALSE
| Amide Hep120220 T240D3_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/92 | 113 | 7010_1_033021062.D | Sample |  | 2021-04-01 12:19 |  | 24.178 | 121894.5758 | FALSE | 123.9631 |  | 123.9631 | 49.6915 | 610.5914 | FALSE | 10.8704 | 6224185.4925 | 56.5912 | FALSE
| Amide Hep120220 T240D4_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/93 | 114 | 7010_1_033021063.D | Sample | 15 | 2021-04-01 12:53 |  | 24.1865 | 152664.7754 | FALSE | 153.8854 |  | 153.8854 | 89.6796 | 604.6658 | FALSE | 10.8704 | 6247719.9279 | 56.6222 | FALSE
| Amide Hep120220 T240D5_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/94 | 115 | 7010_1_033021064.D | Sample |  | 2021-04-01 13:26 |  | 24.1865 | 138382.9851 | FALSE | 136.838 |  | 136.838 | 111.3266 | 603.2422 | FALSE | 10.8704 | 6387251.8875 | 58.8979 | FALSE
| Amide Hep120220 T240D6_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/95 | 116 | 7010_1_033021065.D | Sample | 12 | 2021-04-01 13:59 |  | 24.1865 | 136666.759 | FALSE | 143.1861 |  | 143.1861 | 98.9938 | 600.6554 | FALSE | 10.8704 | 6021869.1892 | 57.9581 | FALSE
| Amide Hep120220 T240D7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/98 | 117 | 7010_1_033021067.D | Sample | 11 | 2021-04-01 15:05 |  | 24.178 | 168104.349 | FALSE | 175.1641 |  | 175.1641 | 90.592 | 601.6773 | FALSE | 10.8704 | 6022115.7937 | 57.9912 | FALSE
| Amide Hep120220 T240D7_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/99 | 117 | 7010_1_033021068.D | Sample | 11 | 2021-04-01 15:39 |  | 24.1695 | 173482.7787 | FALSE | 179.273 |  | 179.273 |  | 601.6896 | FALSE | 10.8704 | 6068137.4872 | 58.3207 | FALSE
| Amide Hep120220 T240D8_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/100 | 118 | 7010_1_033021069.D | Sample | 10 | 2021-04-01 16:12 |  | 24.1695 | 143026.5482 | FALSE | 151.7089 |  | 151.7089 | 113.4237 | 607.5477 | FALSE | 10.8704 | 5939450.8035 | 57.7738 | FALSE
| Amide Hep120220 T240D9_033021 | E:\Shark Tank\Hep\Amides\923 | HC Amides 12/7 1:5 prepped 3/101 | 119 | 7010_1_033021070.D | Sample | 9 | 2021-04-01 16:45 |  | 24.178 | 182171.0493 | FALSE | 191.6037 |  | 191.6037 | 95.9843 | 606.5838 | FALSE | 10.8704 | 5949571.3239 | 59.6113 | FALSE

## Stability_923
Sample | [7 empty] | 923 Results |  |  |  |  | 4NT13C6 (ISTD) Results
| Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | Final Conc. |  | RT | Area | Accuracy | RT | Resp.
| ACN |  |  | 7010_1_031621001.D | MatrixBlank |  | 2021-03-16 10:38 | 0 |  | 24.3896 | 0 |  | 11.0666 | 0
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621002.D | MatrixBlank |  | 2021-03-16 11:11 | 0 |  | 24.5334 | 0 |  | 10.9171 | 3034740.8988
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621003.D | MatrixBlank |  | 2021-03-16 11:45 | 0 |  | 24.5673 | 0 |  | 10.9171 | 2850762.5379
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621004.D | Sample |  | 2021-03-16 12:18 | 0 |  | 24.5757 | 0 |  | 10.9171 | 2644825.5798
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621005.D | MatrixBlank |  | 2021-03-16 12:51 | 0 |  | 24.5588 | 0 |  | 10.9171 | 2887556.6411
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621006.D | Sample |  | 2021-03-16 13:24 | 0 |  | 24.5588 | 0 |  | 10.9171 | 2829167.213
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621007.D | Cal | 15 | 2021-03-16 13:58 | 2062.5121 |  | 24.5503 | 1340428.2035 | 103.1256 | 10.9171 | 3781076.8652
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621008.D | Cal | 15 | 2021-03-16 14:31 | 1960.6723 |  | 24.5504 | 1283927.469 | 98.0336 | 10.9171 | 3836372.3293
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621009.D | Cal | 15 | 2021-03-16 15:04 | 1942.708 |  | 24.5419 | 1263449.977 | 97.1354 | 10.9171 | 3814787.5649
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621010.D | Cal | 15 | 2021-03-16 15:37 | 2065.7418 |  | 24.5334 | 1329550.16 | 103.2871 | 10.9171 | 3743707.0437
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621011.D | Cal | 15 | 2021-03-16 16:11 | 1951.6627 |  | 24.5334 | 1252485.2477 | 97.5831 | 10.9171 | 3762020.4044
| HC_Amide CC 14 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621012.D | Cal | 14 | 2021-03-16 16:44 | 1321.0306 |  | 24.5334 | 832489.4371 | 105.6824 | 10.9078 | 3861574.9417
| HC_Amide CC 13 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621013.D | Cal | 13 | 2021-03-16 17:17 | 714.4068 |  | 24.5334 | 432749.7588 | 91.4733 | 10.9171 | 3884021.5908
| HC_Amide CC 12 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621014.D | Cal | 12 | 2021-03-16 17:50 | 459.2611 |  | 24.5334 | 268250.0373 | 94.1109 | 10.9171 | 3823460.6582
| HC_Amide CC 11 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621015.D | Cal | 11 | 2021-03-16 18:24 | 326.338 |  | 24.5334 | 175504.7195 | 106.9961 | 10.9171 | 3562596.2229
| HC_Amide CC 10 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621016.D | Cal | 10 | 2021-03-16 18:57 | 201.0348 |  | 24.5334 | 115005.6468 | 105.2538 | 10.9171 | 3841091.5489
| HC_Amide CC 9 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621017.D | Cal | 9 | 2021-03-16 19:30 | 118.9905 |  | 24.5334 | 63637.7262 | 99.992 | 10.9078 | 3639188.5513
| HC_Amide CC 8 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621018.D | Cal | 8 | 2021-03-16 20:04 | 63.6042 |  | 24.5334 | 30466.455 | 84.8056 | 10.9171 | 3323239.459
| HC_Amide CC 7 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621019.D | Cal | 7 | 2021-03-16 20:37 | 55.1537 |  | 24.5334 | 30038.2034 | 117.3482 | 10.9078 | 3800055.2559
| HC_Amide CC 6 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621020.D | Cal | 6 | 2021-03-16 21:10 | 32.3369 |  | 24.5081 | 15687.1389 | 111.5067 | 10.9078 | 3483892.0773
| HC_Amide CC 5 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621021.D | Cal | 5 | 2021-03-16 21:44 | 20.0155 |  | 24.5165 | 10094.798 | 111.1972 | 10.9171 | 3779839.2143
| HC_Amide CC 4 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621022.D | Cal | 4 | 2021-03-16 22:17 | 13.5642 |  | 24.5419 | 6787.1748 | 123.3105 | 10.9078 | 3962488.3007
| HC_Amide CC 3 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621023.D | Cal | 3 | 2021-03-16 22:50 | 10.5975 |  | 24.5165 | 4933.2533 | 149.2604 | 10.9078 | 3876163.0483 |  |  |  |  | 923 Stability in Williams' Media
| HC_Amide CC 2 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621024.D | Cal | 2 | 2021-03-16 23:23 | 5.9581 |  | 24.5165 | 2223.5729 | 135.4119 | 10.9078 | 3802153.8368 | Hep Stability Time Course Sample Results |  |  |  |  | % of Time 0
| HC_Amide CC 1 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621025.D | Cal | 1 | 2021-03-16 23:57 | 4.5112 |  | 24.4996 | 1443.4818 | 161.1158 | 10.9078 | 3897229.2312 | Sample ID | Peak Area | Average | Standard Deviation | Time | Replicate | Average | SD | Time (min) | Average Peak Area | Standard Deviation
| Stab_T0A | 3/15/21 Stability |  | 7010_1_031621026.D | Sample | 1 | 2021-03-17 00:30 | 546.2013 | 707.5957 | 24.508 | 334097.9516 |  | 10.9078 | 3975081.6352 | T0 rep 1 | 334097.9516 | 448212.1349 | 99698.1558 | 0 | 74.54% | 1 | 0.22244 | 0 | 448212.1349 | 99698.1558
| Stab_T0B | 3/15/21 Stability |  | 7010_1_031621027.D | Sample | 1 | 2021-03-17 01:03 | 750.4 |  | 24.5165 | 492109.1684 |  | 10.9078 | 4193159.7363 | T0 rep 2 | 492109.1684 |  |  |  | 109.8% |  |  | 5 | 449281.2579 | 14602.1762
| Stab_T0C | 3/15/21 Stability |  | 7010_1_031621028.D | Sample | 1 | 2021-03-17 01:37 | 826.1859 |  | 24.4996 | 518429.2849 |  | 10.9078 | 3988848.4875 | T0 rep 3 | 518429.2849 |  |  |  | 115.7% |  |  | 10 | 400380.5076 | 33745.8463
| Stab_T5A | 3/15/21 Stability |  | 7010_1_031621029.D | Sample | 1 | 2021-03-17 02:10 | 674.3886 | 702.691 | 24.4996 | 437119.4594 |  | 10.9078 | 4169127.058 | T5 rep 1 | 437119.4594 | 449281.2579 | 14602.1762 | 5 | 97.53% | 1.0024 | 0.032579 | 20 | 343655.3914 | 42029.7373
| Stab_T5B | 3/15/21 Stability |  | 7010_1_031621030.D | Sample | 1 | 2021-03-17 02:43 | 702.5621 |  | 24.508 | 445248.229 |  | 10.9078 | 4067340.9205 | T5 rep 2 | 445248.229 |  |  |  | 99.34% |  |  | 30 | 526580.2463 | 197601.102
| Stab_T5C | 3/15/21 Stability |  | 7010_1_031621031.D | Sample | 1 | 2021-03-17 03:17 | 731.1222 |  | 24.4996 | 465476.0853 |  | 10.9078 | 4076910.9199 | T5 rep 3 | 465476.0853 |  |  |  | 103.9% |  |  | 60 | 422907.8995 | 166640.7232
| Stab_T10A | 3/15/21 Stability |  | 7010_1_031621032.D | Sample | 1 | 2021-03-17 03:50 | 591.3768 | 648.1238 | 24.5081 | 365774.4015 |  | 10.9078 | 4004783.8437 | T10 rep 1 | 365774.4015 | 400380.5076 | 33745.8463 | 10 | 81.61% | 0.89328 | 0.07529
| Stab_T10B | 3/15/21 Stability |  | 7010_1_031621033.D | Sample | 1 | 2021-03-17 04:23 | 691.5018 |  | 24.4996 | 433194.6924 |  | 10.9078 | 4024017.6578 | T10 rep 2 | 433194.6924 |  |  |  | 96.65%
| Stab_T10C | 3/15/21 Stability |  | 7010_1_031621034.D | Sample | 1 | 2021-03-17 04:57 | 661.4929 |  | 24.508 | 402172.4289 |  | 10.9078 | 3914576.7753 | T10 rep 3 | 402172.4289 |  |  |  | 89.73%
| HC_Amide CC 10 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621035.D | Cal | 10 | 2021-03-17 05:30 | 188.2778 |  | 24.4827 | 105003.8666 | 98.5747 | 10.9078 | 3750924.0668 | T20 rep 1 | 305373.6074 | 343655.3914 | 42029.7373 | 20 | 68.13% | 0.76672 | 0.093772 | Time (min) | % of Time 0
| HC_Amide CC 5 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621036.D | Cal | 5 | 2021-03-17 06:03 | 18.0635 |  | 24.4911 | 9020.531 | 100.353 | 10.9078 | 3788889.9082 | T20 rep 2 | 336962.6313 |  |  |  | 75.18% |  |  | 0 | 74.54% | 109.8% | 115.7%
| Stab_T20A | 3/15/21 Stability |  | 7010_1_031621037.D | Sample | 1 | 2021-03-17 06:36 | 569.4363 | 589.4414 | 24.4911 | 305373.6074 |  | 10.9078 | 3478467.698 | T20 rep 3 | 388629.9355 |  |  |  | 86.71% |  |  | 5 | 97.53% | 99.34% | 103.9%
| Stab_T20B | 3/15/21 Stability |  | 7010_1_031621038.D | Sample | 1 | 2021-03-17 07:10 | 553.6685 |  | 24.4911 | 336962.6313 |  | 10.9078 | 3952679.3609 | T30 rep 1 | 404142.201 | 526580.2463 | 197601.102 | 30 | 90.17% | 1.1748 | 0.44087 | 10 | 81.61% | 96.65% | 89.73%
| Stab_T20C | 3/15/21 Stability |  | 7010_1_031621039.D | Sample | 1 | 2021-03-17 07:43 | 645.2194 |  | 24.4911 | 388629.9355 |  | 10.9078 | 3883175.112 | T30 rep 2 | 754541.2532 |  |  |  | 168.3% |  |  | 20 | 68.13% | 75.18% | 86.71%
| Stab_T30A | 3/15/21 Stability |  | 7010_1_031621040.D | Sample | 1 | 2021-03-17 08:16 | 640.0215 | 791.6286 | 24.4742 | 404142.201 |  | 10.9078 | 4072651.8038 | T30 rep 3 | 421057.2847 |  |  |  | 93.94% |  |  | 30 | 90.17% | 168.3% | 93.94%
| Stab_T30B | 3/15/21 Stability |  | 7010_1_031621041.D | Sample | 1 | 2021-03-17 08:49 | 1094.2462 |  | 24.4911 | 754541.2532 |  | 10.9078 | 4295927.5781 | T60 rep 1 | 349621.7079 | 422907.8995 | 166640.7232 | 60 | 78% | 0.94354 | 0.37179 | 60 | 78% | 136.9% | 68.15%
| Stab_T30C | 3/15/21 Stability |  | 7010_1_031621042.D | Sample | 1 | 2021-03-17 09:23 | 640.6182 |  | 24.4827 | 421057.2847 |  | 10.9078 | 4238955.8278 | T60 rep 2 | 613632.0818 |  |  |  | 136.9%
| Stab_T06A | 3/15/21 Stability |  | 7010_1_031621043.D | Sample | 1 | 2021-03-17 09:56 | 530.986 | 678.6915 | 24.4742 | 349621.7079 |  | 10.9078 | 4284333.9271 | T60 rep 3 | 305469.9087 |  |  |  | 68.15%
| Stab_T06B | 3/15/21 Stability |  | 7010_1_031621044.D | Sample | 1 | 2021-03-17 10:29 | 1008.2745 |  | 24.4742 | 613632.0818 |  | 10.9078 | 3815842.4912
| Stab_T06C | 3/15/21 Stability |  | 7010_1_031621045.D | Sample | 1 | 2021-03-17 11:02 | 496.8141 |  | 24.4742 | 305469.9087 |  | 10.9078 | 4012124.4191
| HC_Amide CC 7 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621046.D | Cal | 7 | 2021-03-17 11:35 | 46.2601 |  | 24.4826 | 24592.6065 | 98.4257 | 10.9078 | 3739054.1732
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621047.D | MatrixBlank |  | 2021-03-17 12:08 | 0 |  | 24.6857 | 0 |  | 10.9078 | 4548679.7253
Note on J2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: do have final conc here

## Stability_4NT
Sample | [7 empty] | 4NT Results |  |  |  |  | 4NT13C6 (ISTD) Results
| Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | Final Conc. |  | RT | Area | Accuracy | RT | Resp.
| ACN |  |  | 7010_1_031621001.D | MatrixBlank |  | 2021-03-16 10:38 | [5 empty] | 11.0666 | 0
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621002.D | MatrixBlank |  | 2021-03-16 11:11 | 0 |  | 10.8534 | 0 |  | 10.9171 | 3034740.8988
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621003.D | MatrixBlank |  | 2021-03-16 11:45 | 0 |  | 10.8534 | 0 |  | 10.9171 | 2850762.5379
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621004.D | Sample |  | 2021-03-16 12:18 | 0 |  | 10.9188 | 0 |  | 10.9171 | 2644825.5798
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621005.D | MatrixBlank |  | 2021-03-16 12:51 | 0 |  | 10.8441 | 0 |  | 10.9171 | 2887556.6411
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621006.D | Sample |  | 2021-03-16 13:24 | 0 |  | 10.9188 | 0 |  | 10.9171 | 2829167.213
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621007.D | Cal | 15 | 2021-03-16 13:58 | 2004.5677 |  | 10.9188 | 3813432.7839 | 100.2284 | 10.9171 | 3781076.8652
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621008.D | Cal | 15 | 2021-03-16 14:31 | 1975.2267 |  | 10.9189 | 3816529.8286 | 98.7613 | 10.9171 | 3836372.3293
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621009.D | Cal | 15 | 2021-03-16 15:04 | 1978.5415 |  | 10.9188 | 3800979.6399 | 98.9271 | 10.9171 | 3814787.5649
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621010.D | Cal | 15 | 2021-03-16 15:37 | 2000.9083 |  | 10.9188 | 3769338.8724 | 100.0454 | 10.9171 | 3743707.0437
| HC_Amide CC 15 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621011.D | Cal | 15 | 2021-03-16 16:11 | 2013.6975 |  | 10.9188 | 3810261.4566 | 100.6849 | 10.9171 | 3762020.4044
| HC_Amide CC 14 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621012.D | Cal | 14 | 2021-03-16 16:44 | 1336.9265 |  | 10.9095 | 2658901.2854 | 106.9541 | 10.9078 | 3861574.9417
| HC_Amide CC 13 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621013.D | Cal | 13 | 2021-03-16 17:17 | 750.7283 |  | 10.9188 | 1532199.4855 | 96.124 | 10.9171 | 3884021.5908
| HC_Amide CC 12 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621014.D | Cal | 12 | 2021-03-16 17:50 | 440.0265 |  | 10.9188 | 893384.7339 | 90.1694 | 10.9171 | 3823460.6582
| HC_Amide CC 11 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621015.D | Cal | 11 | 2021-03-16 18:24 | 303.1311 |  | 10.9188 | 576090.6307 | 99.3873 | 10.9171 | 3562596.2229
| HC_Amide CC 10 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621016.D | Cal | 10 | 2021-03-16 18:57 | 168.0013 |  | 10.9189 | 345793.5834 | 87.9588 | 10.9171 | 3841091.5489
| HC_Amide CC 9 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621017.D | Cal | 9 | 2021-03-16 19:30 | 117.5952 |  | 10.9095 | 229705.3756 | 98.8195 | 10.9078 | 3639188.5513
| HC_Amide CC 8 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621018.D | Cal | 8 | 2021-03-16 20:04 | 69.9546 |  | 10.9189 | 124980.2677 | 93.2728 | 10.9171 | 3323239.459
| HC_Amide CC 7 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621019.D | Cal | 7 | 2021-03-16 20:37 | 51.5433 |  | 10.9188 | 105363.681 | 109.6667 | 10.9078 | 3800055.2559
| HC_Amide CC 6 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621020.D | Cal | 6 | 2021-03-16 21:10 | 35.0731 |  | 10.9095 | 65766.4157 | 120.9417 | 10.9078 | 3483892.0773
| HC_Amide CC 5 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621021.D | Cal | 5 | 2021-03-16 21:44 | 22.1954 |  | 10.9188 | 45173.8247 | 123.3078 | 10.9171 | 3779839.2143
| HC_Amide CC 4 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621022.D | Cal | 4 | 2021-03-16 22:17 | 13.1884 |  | 10.9095 | 28147.5121 | 119.8945 | 10.9078 | 3962488.3007
| HC_Amide CC 3 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621023.D | Cal | 3 | 2021-03-16 22:50 | 8.4123 |  | 10.9095 | 17565.7 | 118.4831 | 10.9078 | 3876163.0483 |  |  |  |  | 4NT Stability in Williams' Media
| HC_Amide CC 2 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621024.D | Cal | 2 | 2021-03-16 23:23 | 7.8447 |  | 10.9095 | 16068.0865 | 178.2892 | 10.9078 | 3802153.8368 | Hep Stability Time Course Sample Results |  |  |  |  | % of Time 0
| HC_Amide CC 1 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621025.D | Cal | 1 | 2021-03-16 23:57 | 5.8735 |  | 10.9095 | 12332.1119 | 209.7676 | 10.9078 | 3897229.2312 | Sample ID | Peak Area | Average | Standard Deviation | Time | Replicate | Average | SD | Time (min) | Average Peak Area | Standard Deviation
| Stab_T0A | 3/15/21 Stability |  | 7010_1_031621026.D | Sample | 1 | 2021-03-17 00:30 | 747.7695 | 910.6324 | 10.9095 | 1562097.9572 |  | 10.9078 | 3975081.6352 | T0 rep 1 | 1562097.9572 | 1929843.5513 | 318506.4732 | 0 | 80.94% | 1 | 0.16504 | 0 | 1929843.5513 | 318506.4732
| Stab_T0B | 3/15/21 Stability |  | 7010_1_031621027.D | Sample | 1 | 2021-03-17 01:03 | 964.3195 |  | 10.9095 | 2109385.4125 |  | 10.9078 | 4193159.7363 | T0 rep 2 | 2109385.4125 |  |  |  | 109.3% |  |  | 5 | 1800986.3838 | 139799.5587
| Stab_T0C | 3/15/21 Stability |  | 7010_1_031621028.D | Sample | 1 | 2021-03-17 01:37 | 1019.8083 |  | 10.9095 | 2118047.2842 |  | 10.9078 | 3988848.4875 | T0 rep 3 | 2118047.2842 |  |  |  | 109.8% |  |  | 10 | 1583229.5011 | 143134.4102
| Stab_T5A | 3/15/21 Stability |  | 7010_1_031621029.D | Sample | 1 | 2021-03-17 02:10 | 749.451 | 838.3312 | 10.9095 | 1641942.8688 |  | 10.9078 | 4169127.058 | T5 rep 1 | 1641942.8688 | 1800986.3838 | 139799.5587 | 5 | 85.08% | 0.93323 | 0.072441 | 20 | 1341031.193 | 152579.9611
| Stab_T5B | 3/15/21 Stability |  | 7010_1_031621030.D | Sample | 1 | 2021-03-17 02:43 | 872.2575 |  | 10.9095 | 1856575.1749 |  | 10.9078 | 4067340.9205 | T5 rep 2 | 1856575.1749 |  |  |  | 96.2% |  |  | 30 | 1584976.128 | 526121.58
| Stab_T5C | 3/15/21 Stability |  | 7010_1_031621031.D | Sample | 1 | 2021-03-17 03:17 | 893.2853 |  | 10.9095 | 1904441.1076 |  | 10.9078 | 4076910.9199 | T5 rep 3 | 1904441.1076 |  |  |  | 98.68% |  |  | 60 | 912527.0385 | 367601.5895
| Stab_T10A | 3/15/21 Stability |  | 7010_1_031621032.D | Sample | 1 | 2021-03-17 03:50 | 672.3899 | 757.5244 | 10.9095 | 1418741.8341 |  | 10.9078 | 4004783.8437 | T10 rep 1 | 1418741.8341 | 1583229.5011 | 143134.4102 | 10 | 73.52% | 0.82039 | 0.074169
| Stab_T10B | 3/15/21 Stability |  | 7010_1_031621033.D | Sample | 1 | 2021-03-17 04:23 | 781.8545 |  | 10.9095 | 1651497.8222 |  | 10.9078 | 4024017.6578 | T10 rep 2 | 1651497.8222 |  |  |  | 85.58%
| Stab_T10C | 3/15/21 Stability |  | 7010_1_031621034.D | Sample | 1 | 2021-03-17 04:57 | 818.3289 |  | 10.9095 | 1679448.847 |  | 10.9078 | 3914576.7753 | T10 rep 3 | 1679448.847 |  |  |  | 87.03%
| HC_Amide CC 10 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621035.D | Cal | 10 | 2021-03-17 05:30 | 177.1711 |  | 10.9095 | 355998.6592 | 92.7597 | 10.9078 | 3750924.0668 | T20 rep 1 | 1179742.2852 | 1341031.193 | 152579.9611 | 20 | 61.13% | 0.69489 | 0.079063 | Time (min) | % of Time 0
| HC_Amide CC 5 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621036.D | Cal | 5 | 2021-03-17 06:03 | 21.9417 |  | 10.9095 | 44764.6829 | 121.8981 | 10.9078 | 3788889.9082 | T20 rep 2 | 1360274.8272 |  |  |  | 70.49% |  |  | 0 | 80.94% | 109.3% | 109.8%
| Stab_T20A | 3/15/21 Stability |  | 7010_1_031621037.D | Sample | 1 | 2021-03-17 06:36 | 643.0811 | 674.0134 | 10.9095 | 1179742.2852 |  | 10.9078 | 3478467.698 | T20 rep 3 | 1483076.4667 |  |  |  | 76.85% |  |  | 5 | 85.08% | 96.2% | 98.68%
| Stab_T20B | 3/15/21 Stability |  | 7010_1_031621038.D | Sample | 1 | 2021-03-17 07:10 | 652.7449 |  | 10.9095 | 1360274.8272 |  | 10.9078 | 3952679.3609 | T30 rep 1 | 1224896.751 | 1584976.128 | 526121.58 | 30 | 63.47% | 0.8213 | 0.27262 | 10 | 73.52% | 85.58% | 87.03%
| Stab_T20C | 3/15/21 Stability |  | 7010_1_031621039.D | Sample | 1 | 2021-03-17 07:43 | 726.2143 |  | 10.9095 | 1483076.4667 |  | 10.9078 | 3883175.112 | T30 rep 2 | 2188762.2603 |  |  |  | 113.4% |  |  | 20 | 61.13% | 70.49% | 76.85%
| Stab_T30A | 3/15/21 Stability |  | 7010_1_031621040.D | Sample | 1 | 2021-03-17 08:16 | 568.8541 | 715.0074 | 10.9095 | 1224896.751 |  | 10.9078 | 4072651.8038 | T30 rep 3 | 1341269.3727 |  |  |  | 69.5% |  |  | 30 | 63.47% | 113.4% | 69.5%
| Stab_T30B | 3/15/21 Stability |  | 7010_1_031621041.D | Sample | 1 | 2021-03-17 08:49 | 977.097 |  | 10.9095 | 2188762.2603 |  | 10.9078 | 4295927.5781 | T60 rep 1 | 687706.5266 | 912527.0385 | 367601.5895 | 60 | 35.64% | 0.47285 | 0.19048 | 60 | 35.64% | 69.27% | 36.95%
| Stab_T30C | 3/15/21 Stability |  | 7010_1_031621042.D | Sample | 1 | 2021-03-17 09:23 | 599.071 |  | 10.9095 | 1341269.3727 |  | 10.9078 | 4238955.8278 | T60 rep 2 | 1336742.8937 |  |  |  | 69.27%
| Stab_T06A | 3/15/21 Stability |  | 7010_1_031621043.D | Sample | 1 | 2021-03-17 09:56 | 300.8801 | 433.047 | 10.9095 | 687706.5266 |  | 10.9078 | 4284333.9271 | T60 rep 3 | 713131.6952 |  |  |  | 36.95%
| Stab_T06B | 3/15/21 Stability |  | 7010_1_031621044.D | Sample | 1 | 2021-03-17 10:29 | 664.7246 |  | 10.9095 | 1336742.8937 |  | 10.9078 | 3815842.4912
| Stab_T06C | 3/15/21 Stability |  | 7010_1_031621045.D | Sample | 1 | 2021-03-17 11:02 | 333.5362 |  | 10.9095 | 713131.6952 |  | 10.9078 | 4012124.4191
| HC_Amide CC 7 | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621046.D | Cal | 7 | 2021-03-17 11:35 | 52.8956 |  | 10.9095 | 106387.468 | 112.5439 | 10.9078 | 3739054.1732
| Method Spike | HC Amides 12/7 1:5 prepped 3/16 |  | 7010_1_031621047.D | MatrixBlank |  | 2021-03-17 12:08 | 0 |  | 10.9095 | 0 |  | 10.9078 | 4548679.7253
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: also include 3117 stability data even though not including in HepCl report?

## Notes
For sample replicates, I use the first one or report as a concentration +/- sd (%RSD) using data from the replicates (min. of 3 for sd and RSD). It just depends on what your data are for and how you want to present them.
Stability was not ever run on 908, 909 or 916

## ValueList_Helper (hidden)
Sample
Blank
Cal
QC
CC
DoubleBlank
MatrixSpike
MatrixSpikeDup
MatrixBlank
TuneCheck
ResponseCheck